## 2022 catalog
DAI no. | Model | Temp. Range °C | Supplyer | Refrigerant | Charge in gr. | Fyldning tillæg (R14) i gr. | Fyldning i gr. i alt. | GWP value | Climate class | Compressor refrigerator 50/60Hz (50&60Hz) | Compressor Freezer 50/60Hz (50&60Hz) | Flow L/min/ at 10 bar | Flow L/min/ at 9 bar informed by Bolarus | VV længde / dimension | Capillar | Supplier | Volt | Amp | Watt | Hz | Horse power | Målt hos Arctiko Kwh/24h | Oplyst fra producent | Blok Condenser
DAI 0770 | BBR 100 | +4 | BBR 100 | R-290 | 80 |  | 80 | 3.3 |  | TL4CN (Rep. EMT2117U) /EMI50UER |  |  |  |  | ø 0,8 - 3,1m | Bolarus |  |  |  |  | 1/10
DAI 0775-1 | BBR 300 | +4 | BBR 300 | R-290 | 90 |  | 90 | 3.3 |  | EMT 2117U ( Rep. TL4CN)/ EMI50UER |  |  |  |  | ø 1 - 3,6m | Bolarus |  |  |  |  | 1/8
DAI 735 | BBR 500 | +4 | WS - 500 S | R-290 | 110 |  | 110 | 3.3 |  | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus |  |  |  |  | 1/4
DAI 740 | BBR 700 | +4 | WS - 711 S | R-290 | 110 |  | 110 | 3.3 |  | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus |  |  |  |  | 1/4
DAI 745 | BBR 1400 | +4 | WS - 147 S | R-290 | 130 |  | 130 | 3.3 |  | NEK6210U (Rep. NEK2134U) / NLE8.8CNT |  |  |  |  | ø 1,2 - 2,25m | Bolarus |  |  |  |  | 1/2
| BBR 300-D | +4 |  | R-290 | 180 |  | 180 | 3.3 |  | EMT2117U (Rep. TL5CN) / EMI50UER |  |  |  |  | ø 1 - 3,6m | Bolarus
| BBR 500-D | +4 |  | R-290 | 220 |  | 220 | 3.3 | [6 empty] | ø 0,8 - 2,4m | Bolarus
| BBR 700-D | +4 |  | R-290 | 220 |  | 220 | 3.3 |  | EMT2117U, (Rep. EMI70UER) (-20/45) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus
| BBR 1400-D | +4 |  | R-290 | 260 |  | 260 | 3.3 |  | NEK2134U (Rep. NEK6210U ) (-20/45)/ NLE8.8CNT |  |  |  |  | ø 1,2 - 2,25m | Bolarus
DAI 235 | LFF 270 | -30/+4 | S/SN - 300 S | R-290 | R90 + F110 |  | R90 + F110 | 3.3 |  | TL4CN (Rep. EMT2117U) /EMI50UER / | SC12CNX2 (Rep. NEK2134U)/ NT2160U(V), (Rep. SCE15CNLX) |  |  |  | ø 1 - 3,5m, ø 1 - 3,7m | Bolarus |  |  |  |  | 1/10 & 1/8 - 1/3
DAI 265 | LFF 660 | -30/+4 | S/SN - 711 S | R-290 | R90 + F110 |  | R90 + F110 | 3.3 |  | EMT2117U (Rep. TL5CN,EMT6152U) / EMI50UER | SC12CNX2 (Rep. NEK2134U) / NT2160U(V) (Rep. SCE15CNLX) |  |  |  | ø 0,8 - 2,25m, ø 1 - 3m | Bolarus |  |  |  |  | 1/3 & 1/2
DAI 0895 | PR 100 | +1/+10 |  | R-290 | 80 |  | 80 | 3.3 |  | TL4CN, (Rep. EMT2117U) / EMI50UER |  |  |  |  | ø 0,8 - 3,1m | Bolarus |  |  |  |  | 1/8
DAI 0885 | PR 300 | +1/+10 | WS-300 S | R-290 | 90 |  | 90 | 3.3 | 5 | EMT2117U ( Rep. TL4CN) / EMI50UER |  |  |  |  | ø 1 - 3,3m | Bolarus
DAI 850 | PR 500 | +1/+10 | WS - 500 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus |  |  |  |  | 1/3
DAI 860 | PR 700 | +1/+10 | WS - 711 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus |  |  |  |  | 1/3
DAI 860 | PR 700 | +1/+10 | WS - 711 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus
DAI 870 | PR 1400 | +1/+10 | WS - 147 S | R-290 | 130 |  | 130 | 3.3 | 5 | NEK6210U (Rep. NEK2134U) /NLE8.8CNT |  | 23.8 |  |  | ø 1,2 - 2,25 m | Bolarus |  |  |  |  | 1/3
DAI 0280 | LR 100 | +1/+10 | S 100 L | R-290 | 80 |  | 80 | 3.3 |  | TL4CN, (Rep. EMT2117U)/EMI50UER |  |  |  |  | ø 0,8 - 3,1m | Bolarus | 230 | 1.45 |  | 50/60 | 1/8
DAI 230 | LR 270-2 | +1/+10 | S/S - 300 S | R-290 | R90+F90 |  | R90+F90 | 3.3 |  | TL4CN, (Rep. EMT2117U)/EMI50UER |  |  |  |  | ø 1 -3,7m, ø 1 -3,7m | Bolarus |  |  |  |  | 1/8
DAI 220 | LR 300 | +1/+10 | S-300S | R-290 | 90 |  | 90 | 3.3 | 5 | EMT2117U ( Rep. TL4CN)/ EMI50UER |  |  |  |  | ø 1 - 3,3m | Bolarus |  |  |  |  | 1/10 & 1/8
DAI 240 | LR 500 | +1/+10 | S - 500 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus |  |  |  |  | 1/3
DAI 260 | LR 660-2 | +1/+10 | S/S - 711 S |  |  |  |  | 3.3 |  | EMT2117U (Rep. TL5CN, EMT2117U) / EMI50UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus
DAI 250 | LR 700 | +1/+10 | S - 711 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  | 9.5 |  |  | ø 0,8 - 2,4m gl. længde ændret d. 30-01-2020 2380mm 2*0,8mm | Bolarus |  |  |  |  | 1/3
DAI 270 | LR 1400 | +1/+10 | S - 147 S | R-290 | 130 |  | 130 | 3.3 | 5 | NEK6210U (Rep. NEK2134U) / NLE8.8CNT |  |  |  |  | ø 1,2 - 2,25 m | Bolarus |  |  |  |  | 1/2
DAI 6020 | LRE 120 | +2/+8 | Olitrim | R600a | 38 |  | 38 |  |  | HXK80AT | [5 empty] | Olitrim
DAI 6040 | PRE 120 | +2/+8 | Olitrim | R600a | 38 |  | 38 |  |  | HXK80AT | [5 empty] | Olitrim
DAI 6025 | LRE 440 | +2/+8 | Olitrim | R600a | 63 |  | 63 |  |  | NEK2160U | [5 empty] | Olitrim |  |  |  |  | 0.75
DAI 6045 | PRE 440 | +2/+8 | Olitrim | R600a | 63 |  | 63 |  |  | HXK12AT | [5 empty] | Olitrim
DAI 6005 | LFE 110 | -25/-15 | Olitrim | R600a | 37 |  | 37 |  |  | HXK80AT | [5 empty] | Olitrim
| LFE 360 | -25/-15 |  |  | 90 |  | 90
DAI 7010 | LRE 700 | +1/+10 | Ilsa | R290 | 80 |  | 80 | 3.3 |  | EMC3119U | [5 empty] | Ilsa |  |  |  |  | 0.25
DAI 7000 | LFE 700 | -10/-25 | Ilsa | R290 | 120 |  | 120 | 3.3 |  | NEU6168U | [5 empty] | Ilsa |  |  |  |  | 0.75
DAI 7015 | LRE 1400 | +1/+10 | Ilsa | R290 | 90 |  | 90 | 3.3 |  | EMT6165U | [5 empty] | Ilsa |  |  |  |  | 0.33333
DAI 7005 | LFE 1400 | -10/-25 | Ilsa | R290 | 140 |  | 140 | 3.3 |  | NT2180U | [5 empty] | Ilsa |  |  |  |  | 1
| LF 660-2 | -30
DAI 255 | LF 700 | -30 | SN - 711 S | R-290 | 110 |  | 110 | 3.3 | 3 |  | NT2160U (Rep. SC18CNX2) / NT2160U(V), (Rep. SCE18CNLX) |  |  |  | ø 1,2 - 2,25 m | Bolarus | 230 | 4.7 |  | 50 | 5/8 - 3/8
DAI 920 | LAF 700 | -30 |  | R-290 | 150 |  | 150 | 3.3 |  |  | NT2210U / NT2210UV |  |  |  | ø 1,2 - 3,6m
DAI 275 | LF 1400 | -30 | SN - 147 S | R-290 | 160 |  | 160 | 3.3 | 3 |  | NT2210U / NT2210UV |  |  |  | ø 1,5 - 3m | Bolarus |  |  |  |  | 1 1/4
DAI 0285 | LF 100 | -30 | SN 100 L | R-290 | 70 |  | 70 | 3.3 |  |  | SC12CNX2 (Rep. NEK2134U ) / NLE8.8CNT |  |  |  | ø 1 - 3,4m | Bolarus | -20 | 230 | 3.9 |  | 50/60 | 3/8
DAI 225 | LF 300 | -30 | SN - 300 S | R-290 | 100 |  | 100 | 3.3 | 3 |  | SC12CNX2 (Rep. NEK2134U ) / NLE8.8CNT (Rep. SCE15CNLX) |  |  |  | ø 1 - 2,9m | Bolarus |  |  |  | 380 |  | 1/2
DAI 245 | LF 500 | -30 | SN - 500 S | R-290 | 110 |  | 110 | 3.3 | 3 |  | NT2160U(V) (Rep. SCE18CNLX) / NT2160U(V) (Rep. SCE18CNLX) |  |  |  | ø 1,2 - 2,25 m | Bolarus |  |  |  | 580 |  | 3/4
| LTFE 140 | -45 | VT 146 | R290 | 50 |  | 50 | 3.3 |  |  | NEK2134U |  |  |  |  | Vestfrost | -45 |  |  |  |  | 3/8
| LTFE 290 | -45 | VT 306 | R290 | 65 |  | 65 | 3.3 |  |  | NEU2155U |  |  |  |  | Vestfrost | -45 |  |  |  |  | 1/2
| LTFE 370 | -45 | VT 406 | R290 | 70 |  | 70 | 3.3 |  |  | NEU2155U |  |  |  |  | Vestfrost | -45 |  |  |  |  | 1/2
| LTFE 515 | -45 | VT 546 | R290 | 120 |  | 120 | 3.3 |  |  | NEU2155U |  |  |  |  | Vestfrost | [5 empty] | 1/2
DAI 101 | LTF 85 | -60 | VT 76 | Nature-R | 100 g |  | 100 g |  |  |  | NEK2134U |  |  |  |  | Vestfrost | -60 |  |  |  |  | 1/4
| LTF 220 | -60 | VT208 | Nature-R | 137 g |  | 137 g |  |  |  | NEK2134U
DAI 111 | LTF 325 | -60 | VT 306 | Nature-R | 166 g |  | 166 g |  |  |  | NEU2155U |  |  |  |  | Vestfrost | -60 |  |  |  |  | 3/8
DAI 116 | LTF 425 | -60 | VT 407 | Nature-R | 195 g |  | 195 g |  |  |  | NEU2155U |  |  |  |  | Vestfrost | -60 |  |  |  |  | 1/2
DAI 121 | LTF 535 | -60 | VT 547 | Nature-R | 170 g |  | 170 g |  |  |  | SC18CNX.2 |  |  |  |  | Vestfrost | -60 |  |  |  |  | 5/8
DAI 0200-1-S | ULTF 80 | -86 | VT 78 | Nature-R | 117 g |  | 117 g |  |  |  | SC12CNX.2 |  |  |  |  | Vestfrost | -85
DAI 0213 | ULTF 220 | -86 | VT 208 | Nature-R | 108 g |  | 108 g |  |  |  | SC15CNX.2 | 8 |  |  |  | Vestfrost | -85 | 226 | 2.76 | 370 | 50 |  | 6.76 | 7.4
DAI 205 | ULTF 320 | -86 | VT 308 | Nature-R | 143 g |  | 143 g |  |  |  | SC21CNX.2 | 9 |  |  |  | Vestfrost | -85 | 226 | 2.5 | 490 | 50
DAI 210 | ULTF 420 | -86 | VT 408 | Nature-R | 161 g | R600-120gr/R1150-38gr/R50-3gr | 161 g |  |  |  | SC21CNX.2 | 9 |  |  |  | Vestfrost | -85
DAI 1410-1 | MUF 40 | -86 | DTF 35 | EP88 | 139 | 30 | 169 |  |  |  | SC10CNX |  |  |  |  | Hancook | -80 |  |  |  |  | 1/3
DAI 1408-2 | ULUF 450 | -86 | ULUF 450 | EP88 | 270 | 45 | 315 | 5945 |  |  | SC21CNX.2 | Oplyst fra prod. 10,292 | 10.9 | 8500mm / Ø5/Ø8 | 2100mm Ø2,0*0,8 | Bolarus | -85 | 224 | 4.6 | 720 | 50 | 3/4
DAI 1408-3 | ULUF 450 | -86 | ULUF 450 | EP88 | 270 | 45 | 315 | 5945 |  |  | NT2210U |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 224 | 4.6 | 720 | 50 | 1 1/4
DAI 1375 | ULUF 550 | -86 | ULUF 550 | EP 88 | 290 | 68 | 358 |  |  |  | SC21CNX.2 | 9.5 | 10 |  | 2500mm Ø2,0*0,8 | Bolarus | [5 empty] | 3/4
DAI 1416-2 | ULUF 750 | -86 | ULUF 750 | EP88 | (2*270) | (2*45) | 630 |  |  |  | SC21CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus |  |  | 4.6 | 700 | 50 | 3/4
DAI 1390 | ULUF 15 | -86 | ULUF 15 | EP88 | 150 | 23 | 173 |  |  |  | SC12CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 220 | 3.7 | 280 | 50 | 3/8
DAI 1418 | ULUF 65 | -86 | ULUF 65 | EP88 | 171 | 24 | 195 |  |  |  | SC18CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 220 | 2.3 | 496 | 50 | 5/8 | 6.92
DAI 1417 | ULUF 125 | -86 | ULUF 125 | EP88 | 171 | 24 | 195 |  |  |  | SC18CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 220 | 3.3 | 420 | 50 | 5/8
DAI 1800 | ULUF P390GG | -86 | ULUF P390 | GG20 |  |  | 149 | 3.9 | N |  | SC21CNX.2 | 10.9 |  |  | 2100mm Ø2,0*0,8 | Bolarus | -86 |  |  |  |  | 3/4
DAI 1805 | ULUF P500GG | -86 | ULUF P500 | GG20 |  |  | 190 | 3.9 | N |  | SC21CNX.2 | 10.9 |  |  | 2100mm Ø2,0*0,8 | Bolarus | -86 |  |  |  |  | 3/4
DAI 1810 | ULUF P610GG | -86 | ULUF P610 | GG20 |  |  | 219 | 3.9 | N |  | SC21CNX.2 | 10.9 |  |  | 2100mm Ø2,0*0,8 | Bolarus | -86 |  |  |  |  | 3/4
DAI 1815 | ULUF P820GG | -86 | ULUF P820 | GG20 |  |  | 270 | 3.9 | N |  | SC21CNX.2 | 10.9 |  |  | 2100mm Ø2,0*0,8 | Bolarus | -86 |  |  |  |  | 3/4
| ULUF P10 | -86 |  | R600/R1150/R50 |  |  | 76 | 4 |  |  | SC12CNX.2 |  |  |  |  | Vestfrost
| ULUF P50 | -86 |  | R600/R1150/R50 |  |  | 121 | 4 |  |  | SC15CNX.2 |  |  |  |  | Vestfrost
| ULUF P90 | -86 |  | R600/R1150/R50 |  |  | 183 | 4 |  |  | SC18CNX.2 |  |  |  |  | Vestfrost
DAI 0197-S | CRYO 170 | -150 | CLN-1704CWEE | EP88: 450, R14:80g, CH4:30g, AR:35g |  |  | 595 |  |  |  | YS480J2N |  |  |  |  | Hancook | -150 | 400 |  |  | 50
DAI 0198-S | CRYO 230 | -150 | CLN-2304CWEE | EP88: 480, R14:90g, CH4:30g, AR:40g |  |  | 640 |  |  |  | YS480J2N |  |  |  |  | Hancook | -150 | 400 |  |  | 50

## Fyldninger på div enheder
DAI no. | Model | Supplyer | Refrigerant | Charge in gr. | Fyldning tillæg (R14) i gr. | Fyldning i gr. i alt. | GWP value | Climate class | Compressor refrigerator 50/60Hz (50&60Hz) | Compressor Freezer 50/60Hz (50&60Hz) | Flow L/min/ at 10 bar | Flow L/min/ at 9 bar informed by Bolarus | VV længde / dimension | Capillar | Supplier | Temperatur range | Volt | Amp | Watt | Hz | Horse power | Målt hos Arctiko Kwh/24h | Oplyst fra producent | Blok Condenser
DAI 1406 | UPUL 540 | UPUL 540 | R404a | 440 |  | 440 |  |  |  | SC21CL |  |  |  |  | Gram | -40 |  |  |  |  | 3/8
DAI 1404 | UPUL 580 | UPUL 580 | EP88 | 285 | 43 | 328 |  |  |  | SC18CNX | 11.1 |  |  |  | Gram | -85 | 225 | 3.1 | 570 | 50 | 5/8
DAI 1390 | ULUF 15 | ULUF 15 | EP88 | 150 | 23 | 173 |  |  |  | SC12CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 220 | 3.7 | 280 | 50 | 3/8
DAI 1419 | ULUF 60 | ULUF 60 | R404a | 200 |  | 200 |  |  |  | SC18CL |  | 10.3 |  | 2250mm Ø2,0*0,8 | Bolarus | -40 | 220 |  |  |  | 5/8
DAI 1419 | ULUF 60 | ULUF 60 | R404a | [6 empty] | SC18CLX.2 |  | 10.3 |  | 2250mm Ø2,0*0,8
DAI 1418 | ULUF 65 | ULUF 65 | EP88 | 171 | 24 | 195 |  |  |  | SC18CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 220 | 2.3 | 496 | 50 | 5/8 | 6.92
DAI 1395 | ULUF 120 | ULUF 120 | R404a | 200 |  | 200 |  |  |  | SC18CL |  | 10.3 |  | 2250mm Ø2,0*0,8 | Bolarus | -40 | 220 |  |  |  | 5/8
DAI 1395 | ULUF 120 | ULUF 120 | R404a | 150 |  | 150 |  |  |  | SC18CLX.2 |  | 10.3 |  | 2250mm Ø2,0*0,8 | Bolarus
DAI 1417 | ULUF 125 | ULUF 125 | EP88 | 171 | 24 | 195 |  |  |  | SC18CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 220 | 3.3 | 420 | 50 | 5/8
DAI 1407 | ULUF 400 | ULUF 400 | R404a | 400 |  | 400 |  |  |  | SC21CL |  | 11.66 |  | 6800mm Ø2,5*1,0 | Bolarus | -40 |  |  |  |  | 3/8
DAI 1408-2 | ULUF 450 | ULUF 450 | EP88 | 270 | 45 | 315 |  |  |  | SC21CNX.2 | Oplyst fra prod. 10,292 | 10.9 | 8500mm / Ø5/Ø8 | 2100mm Ø2,0*0,8 | Bolarus | -85 | 224 | 4.6 | 720 | 50 | 3/4
DAI 1408-3 | ULUF 450 | ULUF 450 | EP88 | 270 | 45 | 315 |  |  |  | NT2210U |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 224 | 4.6 | 720 | 50 | 1 1/4
DAI 1370 | ULUF 500 | ULUF 500 | R404a | [6 empty] | SC21CL |  |  |  |  | Bolarus | -40 |  |  |  |  | 3/8
DAI 1375 | ULUF 550 | ULUF 550 | EP 88 | 290 | 68 | 358 |  |  |  | SC21CNX.2 | 9.5 | 10 |  | 2500mm Ø2,0*0,8 | Bolarus | [5 empty] | 3/4
DAI 1376-S | ULUF 590 | ULUF 590 | EP 88 | 2*290 | 2*68 | 2*358 |  |  |  | SC21CNX.2 |  | 10 |  | 2500mm Ø2,0*0,8 | Bolarus | [5 empty] | 2*3/4
DAI 1405 | ULUF 490 | ULUF 490 | EP88 | 270 | 45 | 315 |  |  |  | SC21CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 224 | 9.2 | 1440 | 50 | 3/4
DAI 1416 | ULUF 700 | ULUF 700 | R404a | 420 |  | 420 |  |  |  | SC21CL |  |  |  |  | Bolarus | [5 empty] | 3/8
DAI 1416-2 | ULUF 750 | ULUF 750 | EP88 | (2*270) | (2*45) | 630 |  |  |  | SC21CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus |  |  | 4.6 | 700 | 50 | 3/4
DAI 1409 | ULUF 800 | ULUF 800 | R-290 | 140*2 |  | 280 |  |  |  | SC21CL |  |  |  |  | Bolarus | -40 |  |  |  |  | 3/8
DAI 1411 | ULUF 850 | ULUF 850 | EP88 | 270 | 45 | 315 |  |  |  | SC21CNX.2 | Oplyst fra prod. 10,292 | 10.9 | 8500mm / Ø5/Ø8 | 2100mm Ø2,0*0,8 | Bolarus | -85 |  |  |  |  | 3/4
DAI 1411 | ULUF 850 | ULUF 850 | EP88 | 270 | 45 | 315 |  |  |  | NT2210U |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 | 230 |  |  | 50 | 1 1/4
DAI 1415-1 | ULUF 890 | ULUF 890 | EP88 | 540 | 90 | 630 |  |  |  | SC21CNX.2 |  | 10.9 |  | 2100mm Ø2,0*0,8 | Bolarus | -85 |  |  |  |  | 3/4
DAI 1800 | ULUF P390 | ULUF P390 | EP88 | 2*214 | 2*36 | 2x260 |  |  |  | SC21CNX.2 | 10.9 |  |  | 2100mm Ø2,0*0,8 | Bolarus | -86 |  |  |  |  | 3/4
DAI 1805 | ULUF P500 | ULUF P500 | EP88 | 2*262 | 2*43 | 2*305 |  |  |  | SC21CNX.2 | 10.9 |  |  | 2100mm Ø2,0*0,8 | Bolarus | -86 |  |  |  |  | 3/4
DAI 1810 | ULUF P610 | ULUF P610 | EP88 | 2*270 | 2*45 | 2*315 |  |  |  | SC21CNX.2 | 10.9 |  |  | 2100mm Ø2,0*0,8 | Bolarus | -86 |  |  |  |  | 3/4
DAI 1815 | ULUF P820 | ULUF P820 | EP88 | 2*259 | 2*61 | 2*320 |  |  |  | SC21CNX.2 | 10.9 |  |  | 2100mm Ø2,0*0,8 | Bolarus | -86 |  |  |  |  | 3/4
| ULUF P10 |  | R600/R1150/R50 |  |  | 76 | 4 |  |  | SC12CNX.2 |  |  |  |  | Vestfrost
| ULUF P50 |  | R600/R1150/R50 |  |  | 121 | 4 |  |  | SC15CNX.2 |  |  |  |  | Vestfrost
| ULUF P90 |  | R600/R1150/R50 |  |  | 183 | 4 |  |  | SC18CNX.2 |  |  |  |  | Vestfrost
DAI 200 | ULTF 80 | VT 78 | Nature-R | 117 g |  | 117 g |  |  |  | SC12CNX.2 | 7 |  |  |  | Vestfrost | -85
DAI 200 | ULTF 80 | VT 78 | Nature-R | 117 g |  | 117 g |  |  |  | SC12CNX.2 | 7 |  |  |  | Vestfrost | -85 | 224 | 2.1 | 370 | 50
DAI 0200-1-S | ULTF 80 | VT 78 | Nature-R | 117 g |  | 117 g |  |  |  | SC12CNX.2 |  |  |  |  | Vestfrost | -85
DAI 0200-1-S | ULTF 80 | VT 78 | Nature-R | 117 g |  | 117 g |  |  |  | SC12CNX.2 |  |  |  |  | Vestfrost | -85
DAI 0213 | ULTF 220 | VT 208 | Nature-R | 108 g |  | 108 g |  |  |  | SC15CNX.2 | 8 |  |  |  | Vestfrost | -85 | 226 | 2.76 | 370 | 50 |  | 6.76 | 7.4
DAI 0213-1-S | ULTF 220 | VT 208 | Nature-R | 108 g |  | 108 g |  |  |  | SC15CNX.2 |  |  |  |  | Vestfrost | -85
DAI 0213-1-S | ULTF 220 | VT 208 | Nature-R | 108 g |  | 108 g |  |  |  | SC15CNX.2 |  |  |  |  | Vestfrost | -85
DAI 205 | ULTF 320 | VT 308 | Nature-R | 143 g |  | 143 g |  |  |  | SC21CNX.2 | 9 |  |  |  | Vestfrost | -85 | 226 | 2.5 | 490 | 50
DAI 0205-1-S | ULTF 320 | VT 308 | Nature-R | 143 g |  | 143 g |  |  |  | SC21CNX.2 |  |  |  |  | Vestfrost | -85
DAI 205 | ULTF 320 | VT 308 | Nature-R | 143 g |  | 143 g |  |  |  | SC21CNX.2 | 9 |  |  |  | Vestfrost | -85
DAI 205 | ULTF 320 | VT 309 | Nature-R | 143 g |  | 143 g |  |  |  | SC21CNX.2 |  |  |  |  | Vestfrost | -85
DAI 210 | ULTF 420 | VT 408 | Nature-R | 161 g |  | 161 g |  |  |  | SC21CNX.2 | 9 |  |  |  | Vestfrost | -85
DAI 210 | ULTF 420 | VT 408 | Nature-R | 161 g |  | 161 g |  |  |  | SC21CNX.2 | 9 |  |  |  | Vestfrost | -85
DAI 0210-1-S | ULTF 420 | VT 408 | Nature-R | 161 g |  | 161 g |  |  |  | SC21CNX.2 |  |  |  |  | Vestfrost | -85
DAI 210 | ULTF 420 | VT 409 | Nature-R | 161 g |  | 161 g |  |  |  | SC21CNX.2 |  |  |  |  | Vestfrost | -85
DAI 1820 | ULUF P10 | VTS 428 | R600/R1150 | 48/12 | - | 60 |  |  |  | SC12CNX.2 |  |  |  |  | Vestfrost | -86 |  |  |  |  | 3/8
DAI 1825 | ULUF P50 | VTS 438 | R600/R1150/R50 | 78/33,6/1 | - | 112.6 |  |  |  | SC15CNX.2 |  |  |  |  | Vestfrost | -86 |  |  |  |  | 1/2
DAI 1830 | ULUF P90 | VTS 448 | R600/R1150/R50 | 109/38/1 |  | 148 |  |  |  | SC18CNX.2 |  |  |  |  | Vestfrost | -86 |  |  |  |  | 5/8
| LTFG 420 |  | EP88 | 200 | 35 | 235 |  |  |  | SC18CNX |  |  |  |  | Vestfrost | -85 |  |  |  |  | 5/8
DAI 1410 | MUF 40 | DTF 35 | EP88 | 138 | 23 | 161 |  |  |  | NL9CN |  |  |  |  | Hancook | -80 |  |  |  |  | 1/3
DAI 1410-1 | MUF 40 | DTF 35 | EP88 | 138 | 23 | 161 |  |  |  | SC10CNX |  |  |  |  | Hancook | -80 |  |  |  |  | 1/3
DAI 101 | LTF 85 | VT 76 | Nature-R | 100 g |  | 100 g |  |  |  | NEK2134U |  |  |  |  | Vestfrost | -60 |  |  |  |  | 1/4
DAI 101 | LTF 85 | VT 76 | Nature-R | 100 g |  | 100 g |  |  |  | NEK2134U |  |  |  |  | Vestfrost | -60
| LTF 220 | VT208 | Nature-R | 137 g |  | 137 g |  |  |  | NEK2134U
DAI 111 | LTF 325 | VT 306 | Nature-R | 166 g |  | 166 g |  |  |  | NEU2155U |  |  |  |  | Vestfrost | -60 |  |  |  |  | 3/8
DAI 111 | LTF 325 | VT 306 | Nature-R | 166 g |  | 166 g |  |  |  | NEU2155U
DAI 111 | LTF 325 | VT 306 | Nature-R | 166 g |  | 166 g |  |  |  | NEU2155U
DAI 116 | LTF 425 | VT 407 | Nature-R | 195 g |  | 195 g |  |  |  | NEU2155U |  |  |  |  | Vestfrost | -60 |  |  |  |  | 1/2
DAI 116 | LTF 425 | VT 407 | Nature-R | 195 g |  | 195 g |  |  |  | NEU2155U |  |  |  |  | Vestfrost | [5 empty] | 1/2
DAI 121 | LTF 535 | VT 547 | Nature-R | 170 g |  | 170 g |  |  |  | SC18CNX.2 |  |  |  |  | Vestfrost | -60 |  |  |  |  | 5/8
DAI 160 | LTFG 425 | [13 empty] | Vestfrost | -60
| SF 150 | VT 146 USA | EP88 | 170 |  | 170 |  |  |  | SC12CNX.2 |  |  |  |  | Vestfrost | [5 empty] | 3/8
| SF 500 | VT 406 USA | EP88 | 205 |  | 205 |  |  |  | SC15CNX.2 |  |  |  |  | Vestfrost | [5 empty] | 1/2
| SUF 100 | VT 78 USA | EP 88 | 170 | 5 | 175 |  |  |  | SC12CNX.2 |  |  |  |  | Vestfrost | -85 |  |  |  |  | 3/8
| SUF 500 | VT 408 USA | EP 88 | 200 | 35 | 235 |  |  |  | SC18CNX.2 |  |  |  |  | Vestfrost | -85 |  |  |  |  | 5/8
| LTFE 140 | VT 146 | R290 | 50 |  | 50 |  |  |  | NEK2134U |  |  |  |  | Vestfrost | -45 |  |  |  |  | 3/8
| LTFE 290 | VT 306 | R290 | 65 |  | 65 |  |  |  | NEU2155U |  |  |  |  | Vestfrost | -45 |  |  |  |  | 1/2
| LTFE 370 | VT 406 | R290 | 70 |  | 70 |  |  |  | NEU2155U |  |  |  |  | Vestfrost | -45 |  |  |  |  | 1/2
| LTFE 515 | VT 546 | R290 | 120 |  | 120 |  |  |  | NEU2155U |  |  |  |  | Vestfrost | [5 empty] | 1/2
DAI 135 | LTFE 110 | EL11LT | R404a | 110 | [5 empty] | SC12CL |  |  |  |  | Elcold | -45 |  |  |  |  | 3/8
DAI 140 | LTFE 210 | EL21LT | R404a | 180 | [5 empty] | SC15CL |  |  |  |  | Elcold | -45 |  |  |  |  | 1/2
DAI 145 | LTFE 310 | EL31LT | R404a | 200 | [5 empty] | SC18CLX.2 |  |  |  |  | Elcold | -45 |  |  |  |  | 3/4
DAI 150 | LTFE 410 | EL41LT | R404a | 220 | [5 empty] | SC21CL |  |  |  |  | Elcold | -45 |  |  |  |  | 3/8
DAI 155 | LTFE 510 | EL51LT | R404a | 220 | [5 empty] | SC21CL |  |  |  |  | Elcold | -45 |  |  |  |  | 3/8
DAI 0565 | LF 340 | IIRF410 | R134a | 225 |  | 225 |  |  |  | NLY7F |  |  |  |  | Gram |  | 230 | 1.04 | 224 | 50 | 1/3
DAI 0565 | LF 340 | IIRF410 | [7 empty] | NLY9FK |  |  |  |  | Gram
DAI 177 | LCF 100 | PRO 11 | [12 empty] | Elcold
DAI 178 | LCF 200 | PRO 21 | [12 empty] | Elcold
DAI 179 | LCF 300 | PRO 31 | [12 empty] | Elcold
DAI 180 | LCF 400 | PRO 41 | [12 empty] | Elcold
DAI 181 | LCF 500 | PRO 51 | [12 empty] | Elcold
DAI 207 | ULCF 100 | LAB 11 | [12 empty] | Elcold
DAI 208 | ULCF 300 | LAB 31 | [12 empty] | Elcold
DAI 209 | ULCF 400 | LAB 41 | HC (R600 / R1150) ca. 80/20 forhold. |  |  | 230 | [8 empty] | Elcold
DAI 0193 | ULTF 1400 | DW-ZW128 | R 134a 20% R22 20% R 23 20% R 14 30% N2 10% |  |  | 500 eller 600 |  |  |  | TAG 4573T Tecumseh R 22 LRA 61 400 V X 3P/ 50 HZ 440 V X 3P/60HZ |  |  |  |  | Zhongke Meiling |  | 380 | 7 |  | 50/60
DAI0194 | ULTF 1500 | DW-UW258 | R 134a 20% R22 20% R 23 20% R 14 30% N2 10% |  |  | 700 |  |  |  | TAG 4573T Tecumseh R 22 LRA 61 400 V X 3P/ 50 HZ 440 V X 3P/60HZ |  |  |  |  | Zhongke Meiling |  | 380 | 7 |  | 50/60
DAI 0195 | CRYO 170-0 | CLN-1700CWEE | [6 empty] | Uden olie seperator år 2015 | TS370J1N
DAI 0196 | CRYO 230-0 | CLN-2300CWEE | [6 empty] | Uden olie seperator år 2015 | TS370J1N
DAI 0195 | CRYO 170-0 | CLN-1700CWEE |  |  |  | 500 |  |  |  | TS370J1N |  |  |  |  | Hancook | -150
DAI 0196 | CRYO 230-0 | CLN-2300CWEE |  |  |  | 527 |  |  |  | TS370J1N |  |  |  |  | Hancook | -150
DAI 0197-S | CRYO 170 | CLN-1704CWEE | EP88: 450, R14:80g, CH4:30g, AR:35g |  |  | 595 |  |  |  | YS480J2N |  |  |  |  | Hancook | -150 | 400 |  |  | 50
DAI 0198-S | CRYO 230 | CLN-2304CWEE | EP88: 480, R14:90g, CH4:30g, AR:40g |  |  | 640 |  |  |  | YS480J2N |  |  |  |  | Hancook | -150 | 400 |  |  | 50
DAI 6015 | LRE 55 | Bolarus | R134a | 400 |  | 400 |  |  | TL5G | [5 empty] | Bolarus | +4 |  |  |  |  | 0.25
DAI 6035 | PRE 55 | Bolarus | [6 empty] | HTD45AA
DAI 6020 | LRE 120 | Olitrim | R600a | 38 |  | 38 |  |  | HXK80AT | [5 empty] | Olitrim | +4
DAI 6025 | LRE 440 | Olitrim | R600a | 63 |  | 63 |  |  | NEK2160U | [5 empty] | Olitrim | +4 |  |  |  |  | 0.75
DAI 7010 | LRE 700 | Ilsa | R290 | 80 |  | 80 |  |  | EMC3119U | [5 empty] | Ilsa | +4 |  |  |  |  | 0.25
DAI 7015 | LRE 1400 | Ilsa | R290 | 90 |  | 90 |  |  | EMT6165U | [5 empty] | Ilsa | +4 |  |  |  |  | 0.33
DAI 6040 | PRE 120 | Olitrim | R600a | 38 |  | 38 |  |  | HXK80AT | [5 empty] | Olitrim | +4
DAI 6045 | PRE 440 | Olitrim | R600a | 63 |  | 63 |  |  | HXK12AT | [5 empty] | Olitrim | +4
DAI 6005 | LFE 110 | Olitrim | R600a | 37 |  | 37 |  |  | HXK80AT | [5 empty] | Olitrim | -25
DAI 7000 | LFE 700 | Ilsa | R290 | 120 |  | 120 |  |  | NEU2168U | [5 empty] | Ilsa | -25 |  |  |  |  | 0.75
DAI 7005 | LFE 1400 | Ilsa | R290 | 140 |  | 140 |  |  | NT2180U | [5 empty] | Ilsa | -25 |  |  |  |  | 1
DAI 370 | UPLTF 90 | DW-FL90 | R507 | 100 |  | 100 |  |  |  | SC12CL |  |  |  |  | Zhongke Meiling |  | 230 | 2.6 |  | 50 | 3/8
DAI 375 | UPLTF 275 | DW-FL270 | R507 | 350 |  | 350 |  |  |  | SC18CL |  |  |  |  | Zhongke Meiling |  | 230 | 2.7 |  | 50 | 5/8
DAI 360 | UPLTF 330 | GSS3156-20 108 | R404a | 300 |  | 300 |  |  |  | SC15CL |  |  |  |  | Liebherr | -45 |  |  |  |  | 1/2
DAI 0280 | LR 100 | S 100 L | R-290 | 80 |  | 80 | 3.3 |  | TL4CN, (Rep. EMT2117U)/EMI50UER |  |  |  |  | ø 0,8 - 3,1m | Bolarus | 0- +8 | 230 | 1.45 |  | 50/60 | 1/8
DAI 230 | LR 270-2 | S/S - 300 S | R-290 | R90+F90 |  | R90+F90 | 3.3 |  | TL4CN, (Rep. EMT2117U)/EMI50UER |  |  |  |  | ø 1 -3,7m, ø 1 -3,7m | Bolarus | [5 empty] | 1/8
DAI 220 | LR 300 | S-300S | R-290 | 90 |  | 90 | 3.3 | 5 | EMT2117U ( Rep. TL4CN)/ EMI50UER |  |  |  |  | ø 1 - 3,3m | Bolarus | [5 empty] | 1/10 & 1/8
DAI 240 | LR 500 | S - 500 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus | [5 empty] | 1/3
DAI 240 | LR 500 | S - 500 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus |  |  |  | 580 |  | 1/4
DAI 260 | LR 660-2 | S/S - 711 S |  |  |  |  | 3.3 |  | EMT2117U (Rep. TL5CN, EMT2117U) / EMI50UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus
DAI 250 | LR 700 | S - 711 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  | 9.5 |  |  | ø 0,8 - 2,4m gl. længde ændret d. 30-01-2020 2380mm 2*0,8mm | Bolarus | [5 empty] | 1/3
DAI 250 | LR 700 | S - 711 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus
| LR 700-D |  | R-290 | 220 |  | 220 | 3.3 | 5 | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus | +1/+10 | 230 | 1.6 |  |  | 1/3
DAI 0335 | LR 900 | [5 empty] | 3.3 |  | EMT2117U (Rep. TL5CN, EMT2117U) / EMI50UER | [5 empty] | Bolarus
DAI 270 | LR 1400 | S - 147 S | R-290 | 130 |  | 130 | 3.3 | 5 | NEK6210U (Rep. NEK2134U) / NLE8.8CNT |  |  |  |  | ø 1,2 - 2,25 m | Bolarus | [5 empty] | 1/2
DAI 270 | LR 1400 | S - 147 S | R-290 | 130 |  | 130 | 3.3 | 5 | NEK6210U (Rep. NEK2134U) / NLE8.8CNT |  |  |  |  | ø 1,2 - 2,25 m | Bolarus | [5 empty] | 1/10 |  |  | TK1270
DAI 0285 | LF 100 | SN 100 L | R-290 | 70 |  | 70 | 3.3 |  |  | SC12CNX2 (Rep. NEK2134U ) / NLE8.8CNT |  |  |  | ø 1 - 3,4m | Bolarus | -20 | 230 | 3.9 |  | 50/60 | 3/8
DAI 225 | LF 300 | SN - 300 S | R-290 | 100 |  | 100 | 3.3 | 3 |  | SC12CNX2 (Rep. NEK2134U ) / NLE8.8CNT (Rep. SCE15CNLX) |  |  |  | ø 1 - 2,9m | Bolarus |  |  |  | 380 |  | 1/2
DAI 225 | LF 300 | SN - 300 S | R-290 | 100 |  | 100 | 3.3 | 3 |  | SC12CNX2 (Rep. NEK2134U ) / NLE8.8CNT (Rep. SCE15CNLX) |  |  |  | ø 1 - 2,9m | Bolarus |  |  |  | 380 |  | 1/2
DAI 245 | LF 500 | SN - 500 S | R-290 | 110 |  | 110 | 3.3 | 3 |  | NT2160U(V) (Rep. SCE18CNLX) / NT2160U(V) (Rep. SCE18CNLX) |  |  |  | ø 1,2 - 2,25 m | Bolarus |  |  |  | 580 |  | 3/4
DAI 245 | LF 500 | SN - 500 S | R-290 | 110 |  | 110 | 3.3 | 3 |  | NT2160U(V) (Rep. SCE18CNLX) / NT2160U(V) (Rep. SCE18CNLX) |  |  |  | ø 1,2 - 2,25 m | Bolarus
DAI 255 | LF 700 | SN - 711 S | R-290 | 110 |  | 110 | 3.3 | 3 |  | NT2160U (Rep. SC18CNX2) / NT2160U(V), (Rep. SCE18CNLX) |  |  |  | ø 1,2 - 2,25 m | Bolarus | -30 | 230 | 4.7 |  | 50 | 5/8 - 3/8
DAI 0340 | LF 900 | [5 empty] | 3.3 |  |  | NT2178GK |  |  |  |  | Bolarus | -20 | 230 | 3 |  | 50/60 | 3/4
DAI 0340 | LF 900 | [5 empty] | 3.3 |  |  | NT2178GK |  |  |  |  | Bolarus | -20 | 230 | 3 |  | 50/60 | 3/4
DAI 275 | LF 1400 | SN - 147 S | R-290 | 160 |  | 160 | 3.3 | 3 |  | NT2210U / NT2210UV |  |  |  | ø 1,5 - 3m | Bolarus | [5 empty] | 1 1/4
DAI 920 | LAF 700 |  | R-290 | 150 |  | 150 | 3.3 |  |  | NT2210U / NT2210UV |  |  |  | ø 1,2 - 3,6m | Bolarus | -10 til -40 | 230 VC | 6.5 |  | 50 | 1 |  | 9.2
DAI 920 | LAF 700 |  | R-290 | 150 |  | 150 | 3.3 |  |  | NT2210U / NT2210UV |  |  |  | ø 1,2 - 3,6m
DAI 235 | LFF 270 | S/SN - 300 S | R-290 | R90 + F110 |  | R90 + F110 | 3.3 |  | TL4CN (Rep. EMT2117U) /EMI50UER / | SC12CNX2 (Rep. NEK2134U)/ NT2160U(V), (Rep. SCE15CNLX) |  |  |  | ø 1 - 3,5m, ø 1 - 3,7m | Bolarus | [5 empty] | 1/10 & 1/8 - 1/3
DAI 265 | LFF 660 | S/SN - 711 S | R-290 | R90 + F110 |  | R90 + F110 | 3.3 |  | EMT2117U (Rep. TL5CN,EMT6152U) / EMI50UER | SC12CNX2 (Rep. NEK2134U) / NT2160U(V) (Rep. SCE15CNLX) |  |  |  | ø 0,8 - 2,25m, ø 1 - 3m | Bolarus | [5 empty] | 1/3 & 1/2
| LFFG 270 | [7 empty] | TL4CN (Rep. EMT2117U) / SC12CNX2 (Rep. NEK2134U) | SC12CNX2 (Rep. NEK2134U) / EMI50UER / NT2160U(V), (Rep. SCE15CNLX)
| LFFG 660 | [7 empty] | EMT2117U (Rep. TL5CN,EMT6152U) / EMI50UER | / SC12CNX2 (Rep. NEK2134U) / NLE8.8CNT, SCE15CNLX
DAI 0895 | PR 100 |  | R-290 | 80 |  | 80 | 3.3 |  | TL4CN, (Rep. EMT2117U) / EMI50UER |  |  |  |  | ø 0,8 - 3,1m | Bolarus | [5 empty] | 1/8
DAI 0885 | PR 300 | WS-300 S | R-290 | 90 |  | 90 | 3.3 | 5 | EMT2117U ( Rep. TL4CN) / EMI50UER |  |  |  |  | ø 1 - 3,3m | Bolarus
DAI 850 | PR 500 | WS - 500 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus | [5 empty] | 1/3
DAI 860 | PR 700 | WS - 711 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus |  | 220 | 2 | 470 | 50 | 1/3
DAI 860 | PR 700 | WS - 711 S | R-290 | 110 |  | 110 | 3.3 | 5 | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus |  | 220 | 0.72 | 109 | 50
| PR 900 | [5 empty] | 3.3 |  | FR11G
DAI 870 | PR 1400 | WS - 147 S | R-290 | 130 |  | 130 | 3.3 | 5 | NEK6210U (Rep. NEK2134U) /NLE8.8CNT |  |  |  |  | ø 1,2 - 2,25 m | Bolarus | [5 empty] | 1/3
DAI 870 | PR 1400 | WS - 147 S | R-290 | 130 |  | 130 | 3.3 | 5 | NEK6210U (Rep. NEK2134U) /NLE8.8CNT |  | 23.8 |  |  | ø 1,2 - 2,25 m | Bolarus |  |  | 3.5 |  |  | 1/3
DAI 0900 | PF 100 |  | R-290 | 70 |  | 70 | 3.3 |  |  | SC12CNX2 (Rep. NEK2134U ) / NLE8.8CNT |  |  |  | ø 1 - 3,4m | Bolarus | -20 | 230 | 3.9 |  | 50/60 | 3/8
DAI 890 | PF 300 | WSN-300 S | R-290 | 100 |  | 100 | 3.3 | 3 |  | SC12CNX2 (Rep. NEK2134U ) / NLE8.8CNT (Rep. SCE15CNLX) |  |  |  | ø 1 - 2,9m | Bolarus | [5 empty] | 1/2
DAI 855 | PF 500 | WSN - 500 S | R-290 | 110 |  | 110 | 3.3 | 3 |  | NT2160U (Rep. SC18CNX2) / NT2160U(V) (Rep. SCE18CNLX) |  |  |  | ø 1,2 - 2,25 m | Bolarus | -20 | 230 |  | 500 | 50 | 3/8
DAI 855 | PF 500 | WSN - 500 S | R-290 | 110 |  | 110 | 3.3 | 3 |  | NT2160U (Rep. SC18CNX2) / NT2160U(V), (Rep. SCE18CNLX) |  |  |  | ø 1,2 - 2,25 m | Bolarus | [5 empty] | 3/4
DAI 865 | PF 700 | WSN - 711 S | R-290 | 110 |  | 110 | 3.3 | 3 |  | NT2160U (Rep. SC18CNX2) / NT2160U(V), (Rep. SCE18CNLX) |  |  |  | ø 1,2 - 2,25 m | Bolarus |  | 230 |  | 600 | 50 | 3/4
| PF 900 | [5 empty] | 3.3
DAI 875 | PF 1400 | WSN - 147 S | R-290 | 160 |  | 160 | 3.3 | 3 |  | NT2210U / NT2210UV | 36.38 |  |  | ø 1,5 - 3m | Bolarus | [5 empty] | 1/2
DAI 0770 | BBR 100 | BBR 100 | R-290 | 80 |  | 80 | 3.3 |  | TL4CN (Rep. EMT2117U) /EMI50UER |  |  |  |  | ø 0,8 - 3,1m | Bolarus | [5 empty] | 1/10
DAI 0775-1 | BBR 300 | BBR 300 | R-290 | 90 |  | 90 | 3.3 |  | EMT 2117U ( Rep. TL4CN)/ EMI50UER |  |  |  |  | ø 1 - 3,6m | Bolarus |  | 230 | 1.5 |  | 50 | 1/8
DAI 0775-01 | BBR 300 | BBR 300 | R-290 | 90 |  | 90 | 3.3 |  | EMT 2117U ( Rep. TL4CN)/ EMI50UER |  |  |  |  | ø 1 - 3,6m | Bolarus |  |  | 1.5 |  | 50
DAI 735 | BBR 500 | WS - 500 S | R-290 | 110 |  | 110 | 3.3 |  | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus | 0° to +6°C | 230 | 1.7 |  | 50 | 1/4
DAI 735 | BBR 500 | WS - 500 S | R-290 | 110 |  | 110 | 3.3 |  | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus | [5 empty] | 1/3
DAI 740 | BBR 700 | WS - 711 S | R-290 | 110 |  | 110 | 3.3 |  | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus | [5 empty] | 1/4
DAI 740 | BBR 700 | WS - 711 S | R-290 | 110 |  | 110 | 3.3 |  | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus | [5 empty] | 1/3
DAI 740 | BBR 700 | WS - 711 S | R-290 | 110 |  | 110 | 3.3 |  | EMT 2117U (Rep. TL5CN, EMT6152U) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m | Bolarus | [5 empty] | 1/3 |  |  | TK1100
DAI 745 | BBR 1400 | WS - 147 S | R-290 | 130 |  | 130 | 3.3 |  | NEK6210U (Rep. NEK2134U) / NLE8.8CNT |  |  |  |  | ø 1,2 - 2,25m | Bolarus | [5 empty] | 1/2
| BBR 300-D |  | R-290 | 180 |  | 180 | 3.3 |  | EMT2117U (Rep. TL5CN) / EMI50UER |  |  |  |  | ø 1 - 3,6m
| BBR 500-D |  | R-290 | 220 |  | 220 | 3.3 | [6 empty] | ø 0,8 - 2,4m
| BBR 700-D |  | R-290 | 220 |  | 220 | 3.3 |  | EMT2117U, (Rep. EMI70UER) (-20/45) / EMI70UER |  |  |  |  | ø 0,8 - 2,4m
| BBR 1400-D |  | R-290 | 260 |  | 260 | 3.3 |  | NEK2134U (Rep. NEK6210U ) (-20/45)/ NLE8.8CNT |  |  |  |  | ø 1,2 - 2,25m
| VFR700G |  | R134a 180g/R404a 420g/R290 30g/23 180g | [5 empty] | TL5G | SC10CLX |  |  |  |  | Bolarus | [5 empty] | 1/8 & 1/3
DAI 590-1 | PR 170 |  | R134a | 80 |  | 80 |  |  | PL50F / TL5G | [5 empty] | Gram |  | 230 | 0.93 | 106 | 50 | 1/8
DAI 550 | LR 170 | RR200 II | R134a | 50 |  | 50 |  |  | PL50F | [5 empty] | Gram
DAI 550 | LR 170 | RR200 II | R134a | 80 |  | 80 |  |  | TLES3F | [5 empty] | Gram |  | 230 | 0.93 | 106 | 50
DAI 555 | LR 390 | RR 410 | R134a | [5 empty] | TL4G | [5 empty] | Gram | [5 empty] | 1/10
DAI 560-1 | LF 140 | RF200 II | [6 empty] | TLES5,7.FT.3 | [5 empty] | Gram | [5 empty] | 1/4
DAI 560-1 | LF 140 | RF200 II | R134a | 70 |  | 70 |  |  | NL7.3MF | [5 empty] | Gram
| LR 430 |  | R134a | 95 |  | 95 |  |  | NLY9FK | [5 empty] | Gram |  | 230 | 1.16 | 165 | 50
DAI 600 | PR 360 | FKG 410 | R134a | 165 |  | 165 |  |  | FR8.5G | [5 empty] | Vestfrost |  | 220 |  | 320 | 50 | 1/3
DAI 600 | PR 360 | FKG 410 | R134a | 165 |  | 165 |  |  | NL7.3MF | [5 empty] | Vestfrost |  | 220 |  | 320 | 50
| LR 600 |  | R134a | [5 empty] | FR7.5G
DAI 840 | DFS 400 |  | R600a | 48 |  | 48 |  |  |  | ZBH1114CY |  |  |  |  | Westell
DAI 1412 | ULF 390 | DW-86L388 | R404a : 560 gr R290 : 20 gr R508b : 235 gr |  |  | 815 |  |  |  | 2 pcs SC21CL |  |  |  |  | Haier | -40° to -86°C.
DAI 1412 | ULF 390 | DW-86L388 | R134a : 95 gr R404 : 525 gr R290 : 28 gr R23 : 162 gr |  |  | 810 |  |  |  | 2 pcs SC21CL |  |  |  |  | Haier | -40° to -86°C.
DAI 1413 | ULF 650 | DW-86L628 | R134a:180gr R404a:120gr R290:30gr R23:180gr |  |  | 810 |  |  |  | 2 pcs. GS26CLX |  |  |  |  | Haier | -40° to -86°C. | 220-240 | 7 | 1200 | 50
DAI 670 | LAR 70 | HYC-68 | R600a | 22 |  | 22 |  |  | HVY57AA | [5 empty] | Haier
DAI 675 | PHR 70 | HYC-68A | R600a | 22 |  | 22 |  |  | HVY70AA | [5 empty] | Haier
DAI 1012 | PBF 124 | PBF 124
DAI 0173 | VR 110 | MK 304 | R134 | 195 |  | 195 | [8 empty] | Vestfrost |  | 115 |  |  | 60
DAI 0174 | VF 120 | MF 114 | [12 empty] | Vestfrost
DAI 0175 | VF 220 | MF 214 | [12 empty] | Vestfrost
DAI 0176 | VF 320 | MF 314 | [12 empty] | Vestfrost
DAI 7010 | LRE 700 |  | R290 | 80 |  | 80 |  |  | EMC3119U | [5 empty] | Ilsa | +1 / +10 |  |  |  |  | 0.25
DAI 7000 | LFE 700 |  | R290 | 120 |  | 120 |  |  |  | NEU6168U |  |  |  |  | Ilsa | '-10 / -25 |  |  |  |  | 0.75
DAI 7015 | LRE 1400 |  | R290 | 90 |  | 90 |  |  | EMT6165U | [5 empty] | Ilsa | +1 / +10 |  |  |  |  | 0.33333
DAI 7005 | LFE 1400 |  | R290 | 140 |  | 140 |  |  |  | NT2180U |  |  |  |  | Ilsa | '-10 / -25 |  |  |  |  | 1

## Fyldninger ældre enheder
DAI no. | Model | Supplyer | Nationale Lab | Compressor refrigerator | Flow L/min/ at 10 bar | Flow L/min/ at 9 bar informed by Bolarus | VV længde / dimension | Capillar | Compressor Freezer | Refrigerant | Charge in gr. | Fyldning tillæg (R14) i gr. | Fyldning i gr. i alt. | Supplier | Temperatur range | Volt | Amp | Watt | Hz | Horse power | Målt hos Arctiko Kwh/24h | Oplyst fra producent | Blok Condenser
DAI 1406 | UPUL 540 | UPUL 540 | [6 empty] | SC21CL | R404a | 440 |  | 440 | Gram | -40 |  |  |  |  | 3/8
DAI 1404 | UPUL 580 | UPUL 580 | PLPE4586 |  | 11.1 |  |  | Ø2,0*0,8 | SC18CNX | EP88 | 285 | 43 | 328 | Gram | -85 | 225 | 3.1 | 570 | 50 | 5/8
DAI 1390 | ULUF 15 | ULUF 15 |  |  |  | 10.9 |  | 2100mm Ø2,0*0,8 | SC12CNX.2 | EP88 | 150 | 23 | 173 | Bolarus | -85 | 220 | 3.7 | 280 | 50 | 3/8
DAI 1419 | ULUF 60 | ULUF 60 |  |  |  | 10.3 |  | 2250mm Ø2,0*0,8 | SC18CL | R404a | 200 |  | 200 | Bolarus | -40 | 220 |  |  |  | 5/8
DAI 1419 | ULUF 60 | ULUF 60 |  |  |  | 10.3 |  | 2250mm Ø2,0*0,8 | SC18CLX.2 | R404a
DAI 1418 | ULUF 65 | ULUF 65 |  |  |  | 10.9 |  | 2100mm Ø2,0*0,8 | SC18CNX.2 | EP88 | 171 | 24 | 195 | Bolarus | -85 | 220 | 2.3 | 496 | 50 | 5/8 | 6.92
DAI 1395 | ULUF 120 | ULUF 120 |  |  |  | 10.3 |  | 2250mm Ø2,0*0,8 | SC18CL | R404a | 200 |  | 200 | Bolarus | -40 | 220 |  |  |  | 5/8
DAI 1395 | ULUF 120 | ULUF 120 |  |  |  | 10.3 |  | 2250mm Ø2,0*0,8 | SC18CLX.2 | R404a | 150 |  | 150 | Bolarus
DAI 1417 | ULUF 125 | ULUF 125 | PLPE1286EWN |  |  | 10.9 |  | 2100mm Ø2,0*0,8 | SC18CNX.2 | EP88 | 171 | 24 | 195 | Bolarus | -85 | 220 | 3.3 | 420 | 50 | 5/8
DAI 1407 | ULUF 400 | ULUF 400 |  |  |  | 11.66 |  | 6800mm Ø2,5*1,0 | SC21CL | R404a | 400 |  | 400 | Bolarus | -40 |  |  |  |  | 3/8
DAI 1408 | ULUF 450 | ULUF 450 | PLPE4186 |  | Oplyst fra prod. 10,292 | 10.9 | 8500mm / Ø5/Ø8 | 2100mm Ø2,0*0,8 | SC21CNX.2 | EP88 | 270 | 45 | 315 | Bolarus | -85 | 224 | 4.6 | 720 | 50 | 3/4
DAI 1409 | ULUF 450 | ULUF 450 |  |  |  | 10.9 |  | 2100mm Ø2,0*0,8 | NT2210U | EP88 | 270 | 45 | 315 | Bolarus | -85 | 224 | 4.6 | 720 | 50 | 1 1/4
DAI 1370 | ULUF 500 | ULUF 500 | [6 empty] | SC21CL | R404a |  |  |  | Bolarus | -40 |  |  |  |  | 3/8
DAI 1375 | ULUF 550 | ULUF 550 |  |  | 9.5 | 10 |  | 2500mm Ø2,0*0,8 | SC21CNX.2 | EP 88 | 290 | 68 | 358 | Bolarus | [5 empty] | 3/4
DAI 1376-S | ULUF 590 | ULUF 590 |  |  |  | 10 |  | 2500mm Ø2,0*0,8 | SC21CNX.2 | EP 88 | 2*290 | 2*68 | 2*358 | Bolarus | [5 empty] | 2*3/4
DAI 1405 | ULUF 490 | ULUF 490 |  |  |  | 10.9 |  | 2100mm Ø2,0*0,8 | SC21CNX.2 | EP88 | 270 | 45 | 315 | Bolarus | -85 | 224 | 9.2 | 1440 | 50 | 3/4
DAI 1416 | ULUF 700 | ULUF 700 | [6 empty] | SC21CL | R404a | 420 |  | 420 | Bolarus | [5 empty] | 3/8
DAI 1416-2 | ULUF 750 | ULUF 750 |  |  |  | 10.9 |  | 2100mm Ø2,0*0,8 | SC21CNX.2 | EP88 | (2*270) | (2*45) | 630 | Bolarus |  |  | 4.6 | 700 | 50 | 3/4
DAI 1409 | ULUF 800 | ULUF 800 | [6 empty] | SC21CL | R404a | 400 |  | 400 | Bolarus | -40 |  |  |  |  | 3/8
DAI 1411 | ULUF 850 | ULUF 850 |  |  | Oplyst fra prod. 10,292 | 10.9 | 8500mm / Ø5/Ø8 | 2100mm Ø2,0*0,8 | SC21CNX.2 | EP88 | 270 | 45 | 315 | Bolarus | -85 |  |  |  |  | 3/4
DAI 1411 | ULUF 850 | ULUF 850 |  |  |  | 10.9 |  | 2100mm Ø2,0*0,8 | NT2210U | EP88 | 270 | 45 | 315 | Bolarus | -85 | 230 |  |  | 50 | 1 1/4
DAI 1415-1 | ULUF 890 | ULUF 890 |  |  |  | 10.9 |  | 2100mm Ø2,0*0,8 | SC21CNX.2 | EP88 | 540 | 90 | 630 | Bolarus | -85 |  |  |  |  | 3/4
DAI 1800 | ULUF P390 | ULUF P390 | [6 empty] | SC21CNX.2 | EP88 | 2*214 | 2*36 | 2x260 | Bolarus | -86 |  |  |  |  | 3/4
DAI 1805 | ULUF P500 | ULUF P500 | [6 empty] | SC21CNX.2 | EP88 | 2*262 | 2*43 | 2*305 | Bolarus | -86 |  |  |  |  | 3/4
DAI 1810 | ULUF P610 | ULUF P610 | [6 empty] | SC21CNX.2 | EP88 | 2*270 | 2*45 | 2*315 | Bolarus | -86 |  |  |  |  | 3/4
DAI 1815 | ULUF P820 | ULUF P820 | [6 empty] | SC21CNX.2 | EP88 | 2*259 | 2*61 | 2*320 | Bolarus | -86 |  |  |  |  | 3/4
DAI 200 | ULTF 80 | VT 78 |  |  | 7 |  |  |  | SC12CNX | EP88 | 170 | 5 | 175 | Bolarus | -85 |  |  |  |  | 3/8
DAI 200 | ULTF 80 | VT 78 |  |  | 7 |  |  |  | SC12CNX.2 | EP88 | 170 | 5 | 175 | Vestfrost | -85 | 224 | 2.1 | 370 | 50 | 3/8
DAI 0200-1-S | ULTF 80 | VT 78 | [6 empty] | NEK2150U | EP88 | 170 | 5 | 175 | Vestfrost | -85 |  |  |  |  | 1/2
DAI 0200-1-S | ULTF 80 | VT 78 | [6 empty] | NEU2155U | EP88 | 170 | 5 | 175 | Vestfrost | [5 empty] | 3/4
DAI 0213 | ULTF 220 | VT 208 |  |  | 8 |  |  |  | SC15CNX.2 | EP88 | 125 | 10 | 135 | Vestfrost | -85 | 226 | 2.76 | 370 | 50 | 1/2 | 6.76 | 7.4
DAI 0213-1-S | ULTF 220 | VT 208 | [6 empty] | NEK2150U | EP88 | 125 | 10 | 135 | Vestfrost | -85 |  |  |  |  | 1/2
DAI 0213-1-S | ULTF 220 | VT 208 | [6 empty] | NEU2155U | EP88 | 125 | 10 | 135 | Vestfrost | -85
DAI 205 | ULTF 320 | VT 308 | PLTA3386 |  | 9 |  |  |  | SC18CNX | EP88 | 190 | 25 | 215 | Vestfrost | -85 | 226 | 2.5 | 490 | 50 | 5/8
DAI 0205-1-S | ULTF 320 | VT 308 | [6 empty] | NEK2160U | EP88 | 190 | 25 | 215 | Vestfrost | -85 |  |  |  |  | 3/4
DAI 205 | ULTF 320 | VT 308 |  |  | 9 |  |  |  | SC18CNX.2 | EP88 | 190 | 25 | 215 | Vestfrost | -85 |  |  |  |  | 5/8
DAI 205 | ULTF 320 | VT 309 | [6 empty] | SC21CNX.2 | EP88 |  |  |  | Vestfrost | -85 |  |  |  |  | 3/4
DAI 210 | ULTF 420 | VT 408 |  |  | 9 |  |  |  | SC18CNX | EP88 | 200 | 35 | 235 | Vestfrost | -85 |  |  |  |  | 5/8
DAI 210 | ULTF 420 | VT 408 |  |  | 9 |  |  |  | SC18CNX.2 | EP88 | 200 | 35 | 235 | Vestfrost | -85 |  |  |  |  | 5/8
DAI 0210-1-S | ULTF 420 | VT 408 | [6 empty] | NEK2160U | EP88 | 200 | 35 | 235 | Vestfrost | -85 |  |  |  |  | 3/4
DAI 210 | ULTF 420 | VT 409 | [6 empty] | SC21CNX.2 | EP88 |  |  |  | Vestfrost | -85 |  |  |  |  | 3/4
DAI 1820 | ULUF P10 | VTS 428 | [6 empty] | SC12CNX.2 | R600/R1150 | 48/12 | - | 60 | Vestfrost | -86 |  |  |  |  | 3/8
DAI 1825 | ULUF P50 | VTS 438 | [6 empty] | SC15CNX.2 | R600/R1150/R50 | 78/33,6/1 | - | 112.6 | Vestfrost | -86 |  |  |  |  | 1/2
DAI 1830 | ULUF P90 | VTS 448 | [6 empty] | SC18CNX.2 | R600/R1150/R50 | 109/38/1 |  | 148 | Vestfrost | -86 |  |  |  |  | 5/8
| LTFG 420 | [7 empty] | SC18CNX | EP88 | 200 | 35 | 235 | Vestfrost | -85 |  |  |  |  | 5/8
DAI 1410 | MUF 40 | DTF 35 | [6 empty] | NL9CN | EP88 | 138 | 23 | 161 | Hancook | -80 |  |  |  |  | 1/3
DAI 1410-1 | MUF 40 | DTF 35 | [6 empty] | SC10CNX | EP88 | 138 | 23 | 161 | Hancook | -80 |  |  |  |  | 1/3
DAI 100 | LTF 80 | VT 75 | [6 empty] | SC10CL | EP55 | 190 |  | 190 | Vestfrost | -53 |  |  |  |  | 1/3
DAI 105 | LTF 150 | VT 147 | [6 empty] | SC12CL |  |  |  |  | Vestfrost | -53 |  |  |  |  | 3/8
DAI 0126 | LTF 225 | VT 147 -60C | [6 empty] | NEK2134U | HC (R600 / R1150) ca. 80/20 forhold. |  |  | 137 | Vestfrost | -53 |  |  |  |  | 1/2
DAI 110 | LTF 320 | VT 307 | [6 empty] | SC15CL | EP53 | 360 |  | 360 | Vestfrost | -53 |  |  |  |  | 1/2
DAI 115 | LTF 420 | VT 407 | [6 empty] | SC15CLX.2 | EP 53 | 385 |  | 385 | Vestfrost | -53 |  |  |  |  | 1/2
DAI 120 | LTF 530 | VT 536 | [6 empty] | SC18CNX.2 | EP 88 | 190 |  | 190 | Vestfrost | -60 |  |  |  |  | 5/8
DAI 101 | LTF 85 | VT 76 | [6 empty] | NL7CN | HC (R600 / R1150) ca. 80/20 forhold. |  |  | 100 | Vestfrost | -60 |  |  |  |  | 1/4
DAI 101 | LTF 85 | VT 76 | [6 empty] | NEK2134U | HC (R600 / R1150) ca. 80/20 forhold. |  |  | 100 | Vestfrost | -60
DAI 111 | LTF 325 | VT 306 | [6 empty] | SC12CNX.2 | HC (R600 / R1150) ca. 80/20 forhold. |  |  | 170 | Vestfrost | -60 |  |  |  |  | 3/8
DAI 111 | LTF 325 | VT 306 | [6 empty] | NEU2155U
DAI 111 | LTF 325 | VT 306 | [6 empty] | NEK2160U
DAI 116 | LTF 425 | VT 407 | [6 empty] | SC15CNX.2 | HC (R600 / R1150) ca. 80/20 forhold. |  |  | 200 | Vestfrost | -60 |  |  |  |  | 1/2
DAI 116 | LTF 425 | VT 407 | [6 empty] | NEK2150U | HC (R600 / R1150) ca. 80/20 forhold. |  |  | 200 | Vestfrost | [5 empty] | 1/2
DAI 121 | LTF 535 | VT 547 | [6 empty] | SC18CNX.2 | HC (R600 / R1150) ca. 80/20 forhold. |  |  | 180 | Vestfrost | -60 |  |  |  |  | 5/8
DAI 160 | LTFG 425 | [12 empty] | Vestfrost | -60
| SF 150 | VT 146 USA | [6 empty] | SC12CNX.2 | EP88 | 170 |  | 170 | Vestfrost | [5 empty] | 3/8
| SF 500 | VT 406 USA | [6 empty] | SC15CNX.2 | EP88 | 205 |  | 205 | Vestfrost | [5 empty] | 1/2
| SUF 100 | VT 78 USA | [6 empty] | SC12CNX.2 | EP 88 | 170 | 5 | 175 | Vestfrost | -85 |  |  |  |  | 3/8
| SUF 150 | [12 empty] | Vestfrost
| SUF 400 | VT 306 | [6 empty] | SC15CNX | EP88 | 190 |  | 190 | Vestfrost | [5 empty] | 1/2
| SUF 500 | VT 408 USA | [6 empty] | SC18CNX.2 | EP 88 | 200 | 35 | 235 | Vestfrost | -85 |  |  |  |  | 5/8
DAI 1600 | LTFE 140 | [7 empty] | SC12CLX.2 | R404a | 110 |  | 110 | Vestfrost | -45 |  |  |  |  | 3/8
| LTFE 140 | VT 146 | [6 empty] | NEK2134U | R290 | 50 |  | 50 | Vestfrost
DAI 1605 | LTFE 290 | VT 307 -45°C | [6 empty] | SC15CLX.2 | R404a | 132 |  | 132 | Vestfrost | -45 |  |  |  |  | 1/2
| LTFE 290 | VT 306 | [6 empty] | NEU2155U | R290 | 65 |  | 65 | Vestfrost
DAI 1610 | LTFE 370 | VT 406 | [6 empty] | SC15CLX.2 | R404a | 142 |  | 142 | Vestfrost | -45 |  |  |  |  | 1/2
| LTFE 370 | VT 406 | [6 empty] | NEU2155U | R290 | 70 |  | 70 | Vestfrost
DAI 1615 | LTFE 515 | VT 546 | [6 empty] | SC15CLX.2 | R404a | 242 |  | 242 | Vestfrost | [5 empty] | 1/2
| LTFE 515 | VT 546 | [6 empty] | NEU2155U | R290 | 120 |  | 120 | Vestfrost
DAI 135 | LTFE 110 | EL11LT | [6 empty] | SC12CL | R404a | 110 |  |  | Elcold | -45 |  |  |  |  | 3/8
DAI 140 | LTFE 210 | EL21LT | [6 empty] | SC15CL | R404a | 180 |  |  | Elcold | -45 |  |  |  |  | 1/2
DAI 145 | LTFE 310 | EL31LT | [6 empty] | SC18CLX.2 | R404a | 200 |  |  | Elcold | -45 |  |  |  |  | 3/4
DAI 150 | LTFE 410 | EL41LT | [6 empty] | SC21CL | R404a | 220 |  |  | Elcold | -45 |  |  |  |  | 3/8
DAI 155 | LTFE 510 | EL51LT | [6 empty] | SC21CL | R404a | 220 |  |  | Elcold | -45 |  |  |  |  | 3/8
DAI 0565 | LF 340 | IIRF410 | [6 empty] | NLY7F | R134a | 225 |  | 225 | Gram |  | 230 | 1.04 | 224 | 50 | 1/3
DAI 0565 | LF 340 | IIRF410 | [6 empty] | NLY9FK |  |  |  |  | Gram
DAI 177 | LCF 100 | PRO 11 | [11 empty] | Elcold
DAI 178 | LCF 200 | PRO 21 | [11 empty] | Elcold
DAI 179 | LCF 300 | PRO 31 | [11 empty] | Elcold
DAI 180 | LCF 400 | PRO 41 | [11 empty] | Elcold
DAI 181 | LCF 500 | PRO 51 | [11 empty] | Elcold
DAI 207 | ULCF 100 | LAB 11 | [11 empty] | Elcold
DAI 208 | ULCF 300 | LAB 31 | [11 empty] | Elcold
DAI 209 | ULCF 400 | LAB 41 | [7 empty] | HC (R600 / R1150) ca. 80/20 forhold. |  |  | 230 | Elcold
DAI 0193 | ULTF 1400 | DW-ZW128 | [6 empty] | TAG 4573T Tecumseh R 22 LRA 61 400 V X 3P/ 50 HZ 440 V X 3P/60HZ | R 134a 20% R22 20% R 23 20% R 14 30% N2 10% |  |  | 500 eller 600 | Zhongke Meiling |  | 380 | 7 |  | 50/60
DAI0194 | ULTF 1500 | DW-UW258 | [6 empty] | TAG 4573T Tecumseh R 22 LRA 61 400 V X 3P/ 50 HZ 440 V X 3P/60HZ | R 134a 20% R22 20% R 23 20% R 14 30% N2 10% |  |  | 700 | Zhongke Meiling |  | 380 | 7 |  | 50/60
DAI 0195 | CRYO 170-0 | CLN-1700CWEE |  | Uden olie seperator år 2015 |  |  |  |  | TS370J1N | EP88/370g ＋ R14/70g ＋ CH4/30g ＋ Ar/26g
DAI 0196 | CRYO 230-0 | CLN-2300CWEE |  | Uden olie seperator år 2015 |  |  |  |  | TS370J1N | EP88/380g ＋ R14/80g ＋ CH4/32g ＋ Ar/28g
DAI 0195 | CRYO 170-0 | CLN-1700CWEE | [6 empty] | TS370J1N | EP88/370g ＋ R14/70g ＋ CH4/30g ＋ Ar/30g |  |  | 500 | Hancook | -150
DAI 0196 | CRYO 230-0 | CLN-2300CWEE | [6 empty] | TS370J1N | EP88/380g ＋ R14/80g ＋ CH4/32g ＋ Ar/35g |  |  | 527 | Hancook | -150
DAI 0197-S | CRYO 170 | CLN-1704CWEE | [6 empty] | YS480J2N | EP88/450g ＋ R14/80g ＋ CH4/30g ＋ Ar/35g |  |  | 595 | Hancook | -150 | 400 |  |  | 50
DAI 0198-S | CRYO 230 | CLN-2304CWEE | [6 empty] | YS480J2N | EP88/480g ＋ R14/90g ＋ CH4/30g ＋ Ar/40g |  |  | 640 | Hancook | -150 | 400 |  |  | 50
DAI 6015 | LRE 55 | Bolarus |  | TL5G | [5 empty] | R134a | 400 |  | 400 | Bolarus | +4 |  |  |  |  | 0.25
DAI 6035 | PRE 55 | Bolarus
DAI 6020 | LRE 120 | Olitrim |  | HXK80AT | [5 empty] | R600a | 38 |  | 38 | Olitrim | +4
DAI 6025 | LRE 440 | Olitrim |  | NEK2160U | [5 empty] | R600a | 63 |  | 63 | Olitrim | +4 |  |  |  |  | 0.75
DAI 7010 | LRE 700 | Ilsa |  | EMC3119U | [5 empty] | R290 | 80 |  | 80 | Ilsa | +4 |  |  |  |  | 0.25
DAI 7015 | LRE 1400 | Ilsa |  | EMT6165U | [5 empty] | R290 | 90 |  | 90 | Ilsa | +4 |  |  |  |  | 0.33
DAI 6040 | PRE 120 | Olitrim |  | HXK80AT | [5 empty] | R600a | 38 |  | 38 | Olitrim | +4
DAI 6045 | PRE 440 | Olitrim |  | HXK12AT | [5 empty] | R600a | 63 |  | 63 | Olitrim | +4
DAI 6005 | LFE 110 | Olitrim |  | HXK80AT | [5 empty] | R600a | 37 |  | 37 | Olitrim | -25
DAI 7000 | LFE 700 | Ilsa |  | NEU2168U | [5 empty] | R290 | 120 |  | 120 | Ilsa | -25 |  |  |  |  | 0.75
DAI 7005 | LFE 1400 | Ilsa |  | NT2180U | [5 empty] | R290 | 140 |  | 140 | Ilsa | -25 |  |  |  |  | 1
DAI 370 | UPLTF 90 | DW-FL90 | [6 empty] | SC12CL | R507 | 100 |  | 100 | Zhongke Meiling |  | 230 | 2.6 |  | 50 | 3/8
DAI 375 | UPLTF 275 | DW-FL270 | [6 empty] | SC18CL | R507 | 350 |  | 350 | Zhongke Meiling |  | 230 | 2.7 |  | 50 | 5/8
DAI 360 | UPLTF 330 | GSS3156-20 108 | [6 empty] | SC15CL | R404a | 300 |  | 300 | Liebherr | -45 |  |  |  |  | 1/2
DAI 0280 | LR 100 | S 100 L |  | TL5G | [5 empty] | R134a | 250 |  | 250 | Bolarus | 0- +8 | 230 | 1.45 |  | 50/60 | 1/8
DAI 230 | LR 270-2 | S/S - 300 S |  | TL5G | [5 empty] | R134a |  |  |  | Bolarus | [5 empty] | 1/8
DAI 220 | LR 300 | S-300S |  | TL4G, TL5G | [5 empty] | R134a | 190 |  | 190 | Bolarus | [5 empty] | 1/10 & 1/8
DAI 240 | LR 500 | S - 500 S |  | FR8.5G | [5 empty] | R134a | 270 |  | 270 | Bolarus | [5 empty] | 1/3
DAI 240 | LR 500 | S - 500 S |  | NB6152GK, NEK6165GK, EMT6152GK | [5 empty] | R404a | 320 |  | 320 | Bolarus |  |  |  | 580 |  | 1/4
DAI 260 | LR 660-2 | S/S - 711 S | [11 empty] | Bolarus
DAI 250 | LR 700 | S - 711 S |  | FR8.5G | 9.5 |  |  | 2380mm 2*0,8mm |  | R134a | 270 |  | 270 | Bolarus | [5 empty] | 1/3
DAI 250 | LR 700 | S - 711 S |  | NB6252GK | [5 empty] | R404a | 320 |  | 320 | Bolarus
| LR 700-D |  |  | FR8.5G X 2 | [5 empty] | R134a | 420 |  | 420 | Bolarus | +1/+10 | 230 | 1.6 |  |  | 1/3
DAI 0335 | LR 900 |  |  | FR11G | [5 empty] | R134a |  |  |  | Bolarus
DAI 270 | LR 1400 | S - 147 S |  | NEK 6210GK, NEK6213GK, NE2134GK | [5 empty] | R404a | 280 |  | 280 | Bolarus | [5 empty] | 1/2
DAI 270 | LR 1400 | S - 147 S |  | SC10CLX | [5 empty] | R404a | 270 |  | 270 | Bolarus | [5 empty] | 1/10 |  |  | TK1270
DAI 0285 | LF 100 | SN 100 L | [6 empty] | SC12CLX.2 | R404a | 220 |  | 220 | Bolarus | -20 | 230 | 3.9 |  | 50/60 | 3/8
DAI 225 | LF 300 | SN - 300 S | [6 empty] | NEK2150GK | R134a | 280 |  | 280 | Bolarus |  |  |  | 380 |  | 1/2
DAI 225 | LF 300 | SN - 300 S | [6 empty] | NE2134GK | R134a | 320 |  | 320 | Bolarus |  |  |  | 380 |  | 1/2
DAI 245 | LF 500 | SN - 500 S | [6 empty] | T2168GK | R404a | 320 |  | 320 | Bolarus |  |  |  | 580 |  | 3/4
DAI 245 | LF 500 | SN - 500 S | [6 empty] | NT2168GK | R404a | 240 |  | 240 | Bolarus
DAI 255 | LF 700 | SN - 711 S | [6 empty] | SC18CLX.2 SC21CLX.2, T2168GK, NT2168GKV | R404a | 280 |  | 280 | Bolarus | -30 | 230 | 4.7 |  | 50 | 5/8 - 3/8
DAI 0340 | LF 900 | [7 empty] | NT2178GK | R404a | 400 |  | 400 | Bolarus | -20 | 230 | 3 |  | 50/60 | 3/4
DAI 0340 | LF 900 | [7 empty] | NT2178GK | R404a | 300 |  | 300 | Bolarus | -20 | 230 | 3 |  | 50/60 | 3/4
DAI 275 | LF 1400 | SN - 147 S | [6 empty] | NJ2192GK | R404a | 420 |  | 420 | Bolarus | [5 empty] | 1 1/4
DAI 920 | LAF 700 | [7 empty] | GS26CLX | R404a | 1200 |  | 1200 | Bolarus | -10 til -40 | 230 VC | 6.5 |  | 50 | 1 |  | 9.2
DAI 920 | LAF 700 | [7 empty] | GS26CLX | R404a | 384
DAI 235 | LFF 270 | S/SN - 300 S |  | TL4G, TL5G |  |  |  |  | SC10CL, NE2125GK | R134a/R404a | 190/220 |  | 190/220 | Bolarus | [5 empty] | 1/10 & 1/8 - 1/3
DAI 265 | LFF 660 | S/SN - 711 S |  | NB6152GK, NB2116Z, FR8,5G |  |  |  |  | NE2125GK | R134a/R404a | 350 | 290 | 350/290 | Bolarus | [5 empty] | 1/3 & 1/2
DAI 0895 | PR 100 |  |  | TL5G | [5 empty] | R134a | 250 |  | 250 | Bolarus | [5 empty] | 1/8
DAI 0885 | PR 300 | WS-300 S |  | TL5G | [5 empty] | R134a | 190 |  | 190 | Bolarus
DAI 850 | PR 500 | WS - 500 S |  | FR8,5G | [5 empty] | R134a |  |  |  | Bolarus | [5 empty] | 1/3
DAI 860 | PR 700 | WS - 711 S |  | FR8,5G | [5 empty] | R134a | 320 |  | 320 | Bolarus |  | 220 | 2 | 470 | 50 | 1/3
DAI 860 | PR 700 | WS - 711 S |  | NB6152GK | [5 empty] | R404a | 320 |  | 320 | Bolarus |  | 220 | 0.72 | 109 | 50
| PR 900 |  |  | FR11G | [5 empty] | R134a
DAI 870 | PR 1400 | WS - 147 S |  | SC10CLX | [5 empty] | R404a | 270 |  | 270 | Bolarus | [5 empty] | 1/3
DAI 870 | PR 1400 | WS - 147 S |  | UNEK6213GK | 23.8 |  |  |  |  | R404a | 270 |  | 270 | Bolarus |  |  | 3.5 |  |  | 1/3
DAI 0900 | PF 100 | [7 empty] | SC12CLX.2 | R404A | 220 |  | 220 | Bolarus | -20 | 230 | 3.9 |  | 50/60 | 3/8
DAI 890 | PF 300 | WSN-300 S | [6 empty] | NE2134GK | R404a | 320 |  | 320 | Bolarus | [5 empty] | 1/2
DAI 855 | PF 500 | WSN - 500 S | [6 empty] | SC12CL | R404a | 210 |  | 210 | Bolarus | -20 | 230 |  | 500 | 50 | 3/8
DAI 855 | PF 500 | WSN - 500 S | [6 empty] | T2168GK | R404a | 320 |  | 320 | Bolarus | [5 empty] | 3/4
DAI 865 | PF 700 | WSN - 711 S | [6 empty] | T2168GK | R404a | 300 |  | 300 | Bolarus |  | 230 |  | 600 | 50 | 3/4
PF 900
DAI 875 | PF 1400 | WSN - 147 S |  |  | 36.38 |  |  |  | NEK6210GK | R404a | 280 |  | 280 | Bolarus | [5 empty] | 1/2
DAI 0770 | BBR 100 | BBR 100 |  | TL4G | [5 empty] | R134a | 200 |  | 200 | Bolarus | [5 empty] | 1/10
DAI 0775-1 | BBR 300 | BBR 300 |  | TL5G | [5 empty] | R134a | 190 |  | 190 | Bolarus |  | 230 | 1.5 |  | 50 | 1/8
DAI 0775-01 | BBR 300 | BBR 300 | [7 empty] | R134a | 200 |  | 200 | Bolarus |  |  | 1.5 |  | 50
DAI 735 | BBR 500 | WS - 500 S |  | NB6152GK | [5 empty] | R404a | 320 |  | 320 | Bolarus | 0° to +6°C | 230 | 1.7 |  | 50 | 1/4
DAI 735 | BBR 500 | WS - 500 S |  | FR 8.5G | [5 empty] | R134a |  |  |  | Bolarus | [5 empty] | 1/3
DAI 740 | BBR 700 | WS - 711 S |  | NB6152GK | [5 empty] | R404a | 320 |  | 320 | Bolarus | [5 empty] | 1/4
DAI 740 | BBR 700 | WS - 711 S |  | FR 8.5G | [5 empty] | R134a | 270 |  | 270 | Bolarus | [5 empty] | 1/3
DAI 740 | BBR 700 | WS - 711 S |  | FR 8.5G | [5 empty] | R134a | 320 |  | 320 | Bolarus | [5 empty] | 1/3 |  |  | TK1100
DAI 745 | BBR 1400 | WS - 147 S |  | NEK6210GK | [5 empty] | R404a | 270 |  | 270 | Bolarus | [5 empty] | 1/2
BBR 300-D
BBR 500-D
BBR 700-D
BBR 1400-D
| BBR950 | [8 empty] | R134a | 410 |  | 410 | Haier | +2° to +6°C | 220-240 | 3.2 | 930 | 50
| VFR700G |  |  | TL5G |  |  |  |  | SC10CLX | R134a 180g/R404a 420g/R290 30g/23 180g |  |  |  | Bolarus | [5 empty] | 1/8 & 1/3
DAI 590-1 | PR 170 |  |  | PL50F / TL5G | [5 empty] | R134a | 80 |  | 80 | Gram |  | 230 | 0.93 | 106 | 50 | 1/8
DAI 550 | LR 170 | RR200 II |  | PL50F | [5 empty] | R134a | 50 |  | 50 | Gram
DAI 550 | LR 170 | RR200 II |  | TLES3F | [5 empty] | R134a | 80 |  | 80 | Gram |  | 230 | 0.93 | 106 | 50
DAI 555 | LR 390 | RR 410 |  | TL4G | [5 empty] | R134a |  |  |  | Gram | [5 empty] | 1/10
DAI 560-1 | LF 140 | RF200 II |  | TLES5,7.FT.3 | [9 empty] | Gram | [5 empty] | 1/4
DAI 560-1 | LF 140 | RF200 II |  | NL7.3MF | [5 empty] | R134a | 70 |  | 70 | Gram
| LR 430 |  |  | NLY9FK | [5 empty] | R134a | 95 |  | 95 | Gram |  | 230 | 1.16 | 165 | 50
DAI 600 | PR 360 | FKG 410 |  | FR8.5G | [5 empty] | R134a | 165 |  | 165 | Vestfrost |  | 220 |  | 320 | 50 | 1/3
DAI 600 | PR 360 | FKG 410 |  | NL7.3MF | [5 empty] | R134a | 165 |  | 165 | Vestfrost |  | 220 |  | 320 | 50
| LR 600 |  |  | FR7.5G | [5 empty] | R134a
DAI 840 | DFS 400 | [7 empty] | ZBH1114CY | R600a | 48 |  | 48 | Westell
DAI 1412 | ULF 390 | DW-86L388 | [6 empty] | 2 pcs SC21CL | R404a : 560 gr R290 : 20 gr R508b : 235 gr |  |  | 815 | Haier | -40° to -86°C.
DAI 1412 | ULF 390 | DW-86L388 | [6 empty] | 2 pcs SC21CL | R134a : 95 gr R404 : 525 gr R290 : 28 gr R23 : 162 gr |  |  | 810 | Haier | -40° to -86°C.
DAI 1413 | ULF 650 | DW-86L628 | [6 empty] | 2 pcs. GS26CLX | R134a:180gr R404a:120gr R290:30gr R23:180gr |  |  | 810 | Haier | -40° to -86°C. | 220-240 | 7 | 1200 | 50
DAI 670 | LAR 70 | HYC-68 |  | HVY57AA | [5 empty] | R600a | 22 |  | 22 | Haier
DAI 675 | PHR 70 | HYC-68A |  | HVY70AA | [5 empty] | R600a | 22 |  | 22 | Haier
DAI 1012 | PBF 124 | PBF 124
DAI 0173 | VR 110 | MK 304 | [7 empty] | R134 | 195 |  | 195 | Vestfrost |  | 115 |  |  | 60
DAI 0174 | VF 120 | MF 114 | [11 empty] | Vestfrost
DAI 0175 | VF 220 | MF 214 | [11 empty] | Vestfrost
DAI 0176 | VF 320 | MF 314 | [11 empty] | Vestfrost
DAI 7010 | LRE 700 |  |  | EMC3119U | [5 empty] | R290 | 80 |  | 80 | Ilsa | +1 / +10 |  |  |  |  | 0.25
DAI 7000 | LFE 700 | [7 empty] | NEU6168U | R290 | 120 |  | 120 | Ilsa | '-10 / -25 |  |  |  |  | 0.75
DAI 7015 | LRE 1400 |  |  | EMT6165U | [5 empty] | R290 | 90 |  | 90 | Ilsa | +1 / +10 |  |  |  |  | 0.33333
DAI 7005 | LFE 1400 | [7 empty] | NT2180U | R290 | 140 |  | 140 | Ilsa | '-10 / -25 |  |  |  |  | 1
1. | S-300S | LR 300 |  | TL5G, TL4G
2. | S/SN-300S | LR 270-2 |  | chłodnicza V4 TL5G Køl
mroźnicza NL7CLX, NE2125GK, NE2134GK
3. | SN-300S | LF 300 |  | NEK2150GK, NE2134GK
4. | S-500S | LR 500 |  | NB6152GK, NEK6165GK, FR8,5G
5. | WS-500S | PR 500 |  | NB6152GK, NEK6165GK, FR8,5G
6. | SN-500S | LF 500 |  | T2168GK
7. | S-711S | LR 700 |  | FR8.5G
8. | WS-711S | PR 700 |  | NB6152GK, NEK6165GK, FR8,5G
9. | SN-711S | LF 700 |  | T2168GK
10. | S/SN-711S | LFF 660 |  | chłodnicza NB6152GK, NEK6165GK, NB2116Z Køl
mroźnicza NL7CLX, NEK2150GK, NE2125GK, NE2134GK Frys
11. | S-147S | LR 1400 |  | NEK 6210GK, NEK6213GK, NE2134GK
12. | WS-147S | PR 1400 |  | NEK 6210GK, NEK6213GK, NE2134GK
13. | SN-147S | LF 1400 |  | NJ2192GK
ID | Unit name | Leverandør |  | Compressor |  |  |  |  | Refrigerant | Refrigerant charge [kg]
1. | S-300S | LR 300 |  | TL5G |  |  |  |  | R134a | 0.19
2. | S/SN-300S | LFF 270 |  | Refrigerator TL5G |  |  |  |  | R134a | 0.19
|  |  |  | Freezing: NE2125GK |  |  |  |  | R404a | 0.22
3. | SN-300S | LF 300 |  | NEK2150GK |  |  |  |  | R404a | 0.28
|  |  |  | NE2134GK |  |  |  |  | R404a | 0.32
4. | S-500S | LR 500 |  | FR8,5G |  |  |  |  | R134a | 0.27
|  |  |  | NB6152GK |  |  |  |  | R404a | 0.32
5. | WS-500S | PR 500 |  | FR8,5G |  |  |  |  | R134a | 0.27
|  |  |  | NB6152GK |  |  |  |  | R404a | 0.32
6. | SN-500S | LF 500 |  | T2168GK |  |  |  |  | R404a | 0.32
7. | S-711S | LR 700 |  | FR8,5G |  |  |  |  | R134a | 0.27
|  |  |  | NB6152GK |  |  |  |  | R404a | 0.32
8. | WS-711S | PR 700 |  | FR8,5G |  |  |  |  | R134a | 0.27
|  |  |  | NB6152GK |  |  |  |  | R404a | 0.32
9. | SN-711S | LF 700 |  | T2168GK |  |  |  |  | R404a | 0.32
10. | S/SN-711S | LFF 660 |  | - cooling: NB6152GK |  |  |  |  | R404a | 0.32
|  |  |  | FR8,5G |  |  |  |  | R134a | 0.27
|  |  |  | - freezing: NE2125GK |  |  |  |  | R404a | 0.28
|  |  |  | NEK2150GK |  |  |  |  | R404a | 0.29
11. | S-147S | LR 1400 |  | NEK6210GK |  |  |  |  | R404a | 0.28
12. | WS-147S | PR 1400 |  | NEK6210GK |  |  |  |  | R404a | 0.28
13. | SN-147S | LF 1400 |  | NJ2192GK |  |  |  |  | R404a | 0.42
Energiforbrug fået fra Bolarus
kwh/24h
WS-147S dai |  | PR 1400 |  | 8.7
S-500 S /INOX/ |  | LR 500 |  | 2.7
S-711 S /INOX/ OBR/ |  | LR 700 |  | 3.53
S-147 S /INOX//OBR/ |  | LR 1400 |  | 3.07
SN-500 S /INOX/ |  | LF 500 |  | 8.7
SN-711 S /INOX/OBR |  | LF 700 |  | 8.91
SN-147 S /INOX/ |  | LF 1400 |  | 15.1
WS-500 S /INOX/ |  | BBR 500 |  | 5.7
WSN-500 S /INOX/ |  | PF 500 |  | 10.1
WS-711S |  | BBR 700 |  | 6.2
WS-147S |  | PR 1400 |  | 7.7
S/S-300S |  | LR 270-2 |  | 2,4/2,4
S/SN-300S |  | LFF 270 |  | 2,4/7,2
S/S-711S |  | LR 660-2 |  | 2,6/2,6
S/SN-711S |  | LFF 660 |  | 2,6/7,9
|  | Model |  | Bolarus opløsninger | Katalog kwh/24h |  |  |  | Katalog Amp.
|  | BBR 500 |  | 5.7 | 5.7 |  |  |  | 1.2
|  | BBR 700 |  | 6.2 | 3.53 |  |  |  | 1.8
|  | BBR 1400 |  | 8.7 | 8.13 |  |  |  | 3.2
|  | PR 500 |  |  | 5.7 |  |  |  | 1.2
|  | PR 700 |  |  | 3.53 |  |  |  | 1.8
|  | PR 1400 |  |  | 8.13 |  |  |  | 3.2
|  | LR 500 |  | 2.7 | 4.2 |  |  |  | 1
|  | LR 700 |  | 3.53 | 2.95 |  |  |  | 1.6
|  | LR 1400 |  | 3.07 | 8.13 |  |  |  | 2
|  | LF 500 |  | 8.7 | 8.7 |  |  |  | 1.2
|  | LF 700 |  | 8.91 | 9.2 |  |  |  | 4.7
|  | LF 1400 |  | 15.1 | 15.1 |  |  |  | 4.8
|  | PF 500 |  | 10.1 | 10.1 |  |  |  | 1.3
|  | PF 700 |  |  | 9.2 |  |  |  | 5.2
|  | PF 1400 |  |  | 15.1 |  |  |  | 5.2
|  | LR 270-2 |  | 2,4/2,4
|  | LFF 270 |  | 2,4/7,2 | 2,4/8,7 |  |  |  | 1/1
|  | LR 660-2 |  | 2,6/2,6
|  | LFF 660 |  | 2,6/7,9 | 2,95/9,2 |  |  |  | 1,6/4,7

## Beregninger for kapillarrør
Beregninger for kapillarrør ved N2 ved 10 Barrel
Længde af kapillarrør | 1.9946 | meter
Diameter | 0.8 | mm
Flow | 11.1 | L/min
Flow i kapillarrør
Flow | 10.818 | L/min
Længde | 2.1 | meter
Diameter | 0.8 | mm

## Typeskilt info BOL
Standard Units | Voltage [V] | El. current: [A] | Refrigerant | GWP | t CO2 eq | Compresor type | Weight [kg] | Inside temp. [C˚] | Climatic class
LF 100 | 230V 50Hz | 2.8 | R290 | 0 | 0 | SC12CNX.2 | 80 | -10 -30 | 4
LF 100 | 230V 60Hz | 2.8 | R290 | 0 | 0 | NLE8.8CNT | 80 | -10 -30 | 4
LFF 270 | 230V 50Hz / 230V 50Hz | 1,4 / 2,9 | R290 | 0 | 0 | TL4CN / SC12CNX.2 | 130 | +1 +10 / -10 -30 | 4
LFF 270 | 230V 60Hz / 230V 60Hz | 1,4 / 2,9 | R290 | 0 | 0 | EMI50UER / NT2160U | 130 | +1 +10 / -10 -30 | 4
LF 300 | 230V 50Hz | 3.2 | R290 | 0 | 0 | SC12CNX.2 | 110 | -10 -30 | 4
LF 300 | 230V 60Hz | 3.2 | R290 | 0 | 0 | NLE8.8CNT | 110 | -10 -30 | 4
LF 500 | 230V 50Hz | 3.5 | R290 | 0 | 0 | NT2160U | 130 | -10 -30 | 4
LF 500 | 230V 60Hz | 3.5 | R290 | 0 | 0 | NT2160U | 130 | -10 -30 | 4
LF 700 | 230V 50Hz | 3.6 | R290 | 0 | 0 | NT2160U | 135 | -10 -30 | 4
LF 700 | 230V 60Hz | 3.6 | R290 | 0 | 0 | NT2160U | 135 | -10 -30 | 4
LF 700 D | 230V 50Hz | 3.6 | R290 | 0 | 0 | NEU2168U / NEU2168U | 160 | -10 -30 | 4
LFFG 270 | 230V 50Hz / 230V 50Hz | 1,4 / 2,8 | R290 | 0 | 0 | TL4CN / SC12CNX.2 | 140 | +1 +10 / -10 -30 | 4
LFF 270 |  | 2.8 | R290 | 0 | 0 |  |  |  | 4
LF 1400 | 230V 50Hz | 4.5 | R290 | 0 | 0 | NT2210U | 208 | -10 -30 | 4
LF 1400 | 230V 60Hz | 4.5 | R290 | 0 | 0 | NT2210UV | 208 | -10 -30 | 4
LR 100 | 230V 50/60Hz | 1.3 | R290 | 0 | 0 | EMI50UER | 70 | +1 +10 | 4
LR 270-2 | 230V 50Hz / 230V 50Hz | 1,4 / 1,4 | R290 | 0 | 0 | TL4CN / TL4CN | 120 | +1 +10 / +1 +10 | 4
LR 270 |  | 1.4 | R290 | 0 | 0 |  |  |  | 4
LR 300 | 230V 50Hz | 1.7 | R290 | 0 | 0 | TL4CN | 100 | +1 +10 | 4
LR 300 | 230V 60Hz | 1.7 | R290 | 0 | 0 | EMI50UER | 100 | +1 +10 | 4
LR 500 | 230V 50Hz | 1.9 | R290 | 0 | 0 | EMT2117U | 115 | +1 +10 | 4
LR 500 | 230V 60Hz | 1.9 | R290 | 0 | 0 | EMI70UER | 115 | +1 +10 | 4
LR 700 | 230V 50Hz | 1.9 | R290 | 0 | 0 | EMT2117U | 120 | +1 +10 | 4
LR 700 | 230V 60Hz | 1.9 | R290 | 0 | 0 | EMI70UER | 120 | +1 +10 | 4
LR 700 D | 230V 50/60Hz | 1.9 | R290 | 0 | 0 | EMI70UER / EMI70UER | 165 | +1 +10 | 4
LR 1400 | 230V 50Hz | 2.5 | R290 | 0 | 0 | NEK6210U | 170 | +1 +10 | 4
LR 1400 | 230V 60Hz | 2.5 | R290 | 0 | 0 | NLE8.8CNT | 170 | +1 +10 | 4
LFFG 660 | 230V 50Hz / 230V 50Hz | 1,7/2,9 | R290 | 0 | 0 | EMT2117U / SC12CNX.2 | 167 | +1 +10 / -10 -30 | 4
LFF 660 | 230V 50Hz / 230V 50Hz | 1,7/2,9 | R290 | 0 | 0 | EMT2117U / SC12CNX.2 | 160 | +1 +10 / -10 -30 | 4
LF 660-2 | 230V 50Hz / 230V 50Hz | 2,9 / 2,9 | R290 | 0 | 0 | SC12CNX.2 / SC12CNX.2 | 165 | -10 -30 / -10 -30 | 4
PR 100 | 230V 50Hz | 1.4 | R290 | 0 | 0 | TL4CN | 85 | +1 +10 | 4
PR 100 | 230V 60Hz | 1.4 | R290 | 0 | 0 | EMI50UER | 85 | +1 +10 | 4
PR 300 | 230V 50Hz | 1.8 | R290 | 0 | 0 | TL4CN | 110 | +1 +10 | 4
PR 300 | 230V 60Hz | 1.8 | R290 | 0 | 0 | EMI50UER | 110 | +1 +10 | 4
PR 500 | 230V 50Hz | 1.9 | R290 | 0 | 0 | EMT2117U | 135 | +1 +10 | 4
PR 500 | 230V 60Hz | 1.9 | R290 | 0 | 0 | EMI70UER | 135 | +1 +10 | 4
PR 700 | 230V 50Hz | 2 | R290 | 0 | 0 | EMT2117U | 150 | +1 +10 | 4
PR 700 | 230V 60Hz | 2 | R290 | 0 | 0 | EMI70UER | 150 | +1 +10 | 4
PR 1400 | 230V 50Hz | 2.6 | R290 | 0 | 0 | NEK6210U | 207 | +1 +10 | 4
PR 1400 | 230V 60Hz | 2.6 | R290 | 0 | 0 | NLE8.8CNT | 207 | +1 +10 | 4
PR 1400 D | 230V 50Hz | 2.7 | R290 | 0 | 0 | NEK2134U / NEK2134U | 230 | +1 +10 | 4
PF 100 | 230V 50Hz | 2.9 | R290 | 0 | 0 | SC12CNX.2 | 90 | -10 -23 | 4
PF 300 | 230V 50Hz | 3.4 | R290 | 0 | 0 | SC12CNX.2 | 115 | -10 -23 | 4
PF 300 | 230V 60Hz | 3.4 | R290 | 0 | 0 | NLE8.8CNT | 115 | -10 -23 | 4
PF 500 | 230V 50Hz | 3.5 | R290 | 0 | 0 | NT2160U | 140 | -10 -23 | 4
PF 500 | 230V 60Hz | 3.5 | R290 | 0 | 0 | NT2160U | 140 | -10 -23 | 4
PF 700 | 230V 50Hz | 3.6 | R290 | 0 | 0 | NT2160U | 160 | -10 -23 | 4
PF 700 | 230V 60Hz | 3.6 | R290 | 0 | 0 | NT2160U | 160 | -10 -23 | 4
PF 1400 | 230V 50Hz | 7 | R290 | 0 | 0 | NT2210U | 220 | -10 -23 | 4
LAF 700 | 230V 50Hz | 4 | R290 | 0 | 0 | NT2210U | 151 | -10 -40 | 4
LAF 700 | 230V 60Hz | 4 | R290 | 0 | 0 | NT2210UV | 151 | -10 -40 | 4
BBR Units | Voltage [V] | El. current: [A] | Refrigerant | GWP | t CO2 eq | Compresor type | Weight [kg] | Inside temp. [C˚] | Climatic class
BBR 100 | 230V 50Hz | 1.4 | R290 | 0 | 0 | TL4CN | 71 | +4 | 4
BBR 100 | 230V 60Hz | 1.4 | R290 | 0 | 0 | EMI50UER | 71 | +4 | 4
BBR 100 D | 230V 50Hz | 1,4 / 1,4 | R290 | 0 | 0 | EMI50UER / EMI50UER | 85 | +4 | 4
BBR 300 | 230V 50Hz | 1.7 | R290 | 0 | 0 | TL4CN | 105 | +4 | 4
BBR 300 | 230V 60Hz | 1.7 | R290 | 0 | 0 | EMI50UER | 105 | +4 | 4
BBR 300 D | 230V 50Hz | 1,7 / 1,7 | R290 | 0 | 0 | EMT2117U / EMT2117U | 120 | +4 | 4
BBR 500 | 230V 50/60Hz | 2 | R290 | 0 | 0 | EMI70UER | 120 | +4 | 4
BBR 500 D | 230V 50Hz | 2 / 2 | R290 | 0 | 0 | EMI70UER / EMI70UER | 130 | +4 | 4
BBR 700 | 230V 50/60Hz | 2.1 | R290 | 0 | 0 | EMI70UER / EMI70UER | 150 | +4 | 4
BBR 700 D | 230V 50/60Hz | 2,1 / 2,1 | R290 | 0 | 0 | EMI70UER / EMI70UER | 160 | +4 | 4
BBR 1400 | 230V 50Hz | 2.6 | R290 | 0 | 0 | NEK6210U | 210 | +4 | 4
BBR 1400 | 230V 60Hz | 2.6 | R290 | 0 | 0 | NLE8.8CNT | 210 | +4 | 4
BBR 1400 D | 230V 50Hz | 2,6 / 2,6 | R290 | 0 | 0 | NEK6210U / NEK6210U | 240 | +4 | 4
ULUF Units | Voltage [V] | El. current: [A] | Refrigerant | GWP | t CO2 eq | Compresor type | Weight [kg] | Inside temp. [C˚] | Climatic class | ULUF P390 | 230V - 50/60Hz - 8,5A | 2 x SC21CNX.2 | Refrigerant: 2 x 250g EP88 (R14/23/116/290/600a) | GWP 5945 | 3,0 t CO2 eq | Weight: 215 kg
ULUF 15 | 230V 50/60Hz | 2.8 | EP88 | GWP 5945 | 1,0 t CO2 eq | SC12CNX.2 | 38 | -40 -86 |  | ULUF P500 | 230V - 50/60Hz - 8,9A | 2 x SC21CNX.2 | Refrigerant: 2 x 305g EP88 (R14/23/116/290/600a) | GWP 5945 | 3,6 t CO2 eq | Weight: 231 kg
ULUF 65 | 230V 50/60Hz | 3.5 | EP88 | GWP 5945 | 1,2 t CO2 eq | SC18CNX.2 | 77 | -40 -86 |  | ULUF P610 | 230V - 50/60Hz - ??A | 2 x SC21CNX.2 | Refrigerant: 2 x 315g EP88 (R14/23/116/290/600a) | GWP 5945 | 3,7 t CO2 eq | Weight: 245 kg
ULUF 125 | 230V 50/60Hz | 3.3 | EP88 | GWP 5945 | 1,2 t CO2 eq | SC18CNX.2 | 98 | -40 -86 |  | ULUF P820 | 230V - 50/60Hz - 10A | 2 x SC21CNX.2 | Refrigerant: 2 x 320g EP88 (R14/23/116/290/600a) | GWP 5945 | 3,8 t CO2 eq | Weight: 273 kg
ULUF 450 | 230V 50/60Hz | 4.4 | EP88 | GWP 5945 | 1,9 t CO2 eq | SC21CNX.2 | 198 | -40 -86
ULUF 450-2M | 230V 50/60Hz | 4.4 | EP88 | GWP 5945 | 1,9 t CO2 eq | SC21CNX.2 | 175 | -40 -86
ULUF 490 | 230V 50/60Hz | 8.3 | EP88 | GWP 5945 | 3,7 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 198 | -40 -86 |  | ULUF 15 | 230V - 50/60Hz - 2,8A | SC12CNX.2 | Refrigerant: 173g EP88 (R14/23/116/290/600a) | GWP 5945 | 1,0 t CO2 eq | Weight: 38 kg
ULUF 490-2M | 230V 50/60Hz | 8.3 | EP88 | GWP 5945 | 3,7 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 175 | -40 -86 |  | ULUF 65 | 230V - 50/60Hz - 3,5A | SC18CNX.2 | Refrigerant: 195g EP88 (R14/23/116/290/600a) | GWP 5945 | 1,2 t CO2 eq | Weight: 77 kg
ULUF 550 | 230V 50/60Hz | 4.2 | EP88 | GWP 5945 | 2,1 t CO2 eq | SC21CNX.2 | 215 | -40 -86 |  | ULUF 125 | 230V - 50/60Hz - 3,3A | SC18CNX.2 | Refrigerant: 195g EP88 (R14/23/116/290/600a) | GWP 5945 | 1,2 t CO2 eq | Weight: 98 kg
ULUF 550-2M | 230V 50/60Hz | 4.2 | EP88 | GWP 5945 | 2,1 t CO2 eq | SC21CNX.2 | 203 | -40 -86 |  | ULUF 450 | 230V - 50/60Hz - 4,4A | SC21CNX.2 | Refrigerant: 315g EP88 (R14/23/116/290/600a) | GWP 5945 | 1,9 t CO2 eq | Weight: 198 kg
ULUF 590 | 230V 50/60Hz | 8.2 | EP88 | GWP 5945 | 4,3 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 215 | -40 -86 |  | ULUF 450-2M | 230V - 50/60Hz - 4,4A | SC21CNX.2 | Refrigerant: 315g EP88 (R14/23/116/290/600a) | GWP 5945 | 1,9 t CO2 eq | Weight: 175 kg
ULUF 590-2M | 230V 50/60Hz | 8.2 | EP88 | GWP 5945 | 4,3 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 203 | -40 -86
ULUF 750 | 230V 50/60Hz | 8.7 | EP88 | GWP 5945 | 3,7 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 254 | -40 -86 |  | ULUF 490 | 230V - 50/60Hz - 8,3A | 2 x SC21CNX.2 | Refrigerant: 315g 2 x EP88 (R14/23/116/290/600a) | GWP 5945 | 3,7 t CO2 eq | Weight: 198 kg
ULUF 750-2M | 230V 50/60Hz | 8.7 | EP88 | GWP 5945 | 3,7 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 242 | -40 -86 |  | ULUF 490-2M | 230V - 50/60Hz - 8,3A | 2 x SC21CNX.2 | Refrigerant: 315g 2 x EP88 (R14/23/116/290/600a) | GWP 5945 | 3,7 t CO2 eq | Weight: 175 kg
ULUF 850 | 230V 50/60Hz | 8.3 | EP88 | GWP 5945 | 1,9 t CO2 eq | SC21CNX.2 | 401 | -40 -86 |  | ULUF 550 | 230V - 50/60Hz - 4,2A | SC21CNX.2 | Refrigerant: 358g EP88 (R14/23/116/290/600a) | GWP 5945 | 2,1 t CO2 eq | Weight: 215 kg
ULUF 850-2M | 230V 50/60Hz | 8.3 | EP88 | GWP 5945 | 1,9 t CO2 eq | SC21CNX.2 | 355 | -40 -86 |  | ULUF 550-2M | 230V - 50/60Hz - 4,2A | SC21CNX.2 | Refrigerant: 358g EP88 (R14/23/116/290/600a) | GWP 5945 | 2,1 t CO2 eq | Weight: 203 kg
ULUF 890 | 230V 50/60Hz | 17 | EP88 | GWP 5945 | 3,7 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 490 | -40 -86 |  | ULUF 590 | 230V - 50/60Hz - 8,2A | 2 x SC21CNX.2 | Refrigerant: 358g 2 x EP88 (R14/23/116/290/600a) | GWP 5945 | 4,3 t CO2 eq | Weight: 215 kg
ULUF 890-2M | 230V 50/60Hz | 17 | EP88 | GWP 5945 | 3,7 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 440 | -40 -86 |  | ULUF 590-2M | 230V - 50/60Hz - 8,2A | 2 x SC21CNX.2 | Refrigerant: 358g 2 x EP88 (R14/23/116/290/600a) | GWP 5945 | 4,3 t CO2 eq | Weight: 203 kg
[10 empty] | ULUF 750 | 230V - 50/60Hz - 8,7A | 2 x SC21CNX.2 | Refrigerant: 315g 2 x EP88 (R14/23/116/290/600a) | GWP 5945 | 3,7 t CO2 eq | Weight: 254 kg
[10 empty] | ULUF 750-2M | 230V - 50/60Hz - 8,7A | 2 x SC21CNX.2 | Refrigerant: 315g 2 x EP88 (R14/23/116/290/600a) | GWP 5945 | 3,7 t CO2 eq | Weight: 242 kg
ULUF P Units | Voltage [V] | El. current: [A] | Refrigerant | GWP | t CO2 eq | Compresor type | Weight [kg] | Inside temp. [C˚] | Climatic class | ULUF 850 | 230V - 50/60Hz - 8,3A | SC21CNX.2 | Refrigerant: 315g EP88 (R14/23/116/290/600a) | GWP 5945 | 1,9 t CO2 eq | Weight: 401 kg
ULUF P390 | 230V 50/60Hz | 8.5 | EP88 | GWP 5945 | 3,0 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 215 | -40 -90 |  | ULUF 850-2M | 230V - 50/60Hz - 8,3A | SC21CNX.2 | Refrigerant: 315g EP88 (R14/23/116/290/600a) | GWP 5945 | 1,9 t CO2 eq | Weight: 355 kg
ULUF P 500 | 230V 50/60Hz | 8.9 | EP88 | GWP 5945 | 3,6 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 231 | -40 -90 |  | ULUF 890 | 230V - 50/60Hz - 17A | 2 x SC21CNX.2 | Refrigerant: 315g 2 x EP88 (R14/23/116/290/600a) | GWP 5945 | 3,7 t CO2 eq | Weight: 490 kg
ULUF P 610 | 230V 50/60Hz | 9.3 | EP88 | GWP 5945 | 3,7 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 245 | -40 -90 |  | ULUF 890-2M | 230V - 50/60Hz - 17A | 2 x SC21CNX.2 | Refrigerant: 315g 2 x EP88 (R14/23/116/290/600a) | GWP 5945 | 3,7 t CO2 eq | Weight: 440 kg
ULUF P 820 | 230V 50/60Hz | 10 | EP88 | GWP 5945 | 3,8 t CO2 eq | SC21CNX.2 / SC21CNX.2 | 273 | -40 -90
ULUF P 390 GG | 230V 50/60Hz | 8.6 | Green Gas | GWP3,9 |  | SC21CNX.2 / SC21CNX.2 | 212 | -40 -86 | N
ULUF P 500 GG | 230V 50/60Hz | 8.7 | Green Gas | GWP3,9 |  | SC21CNX.2 / SC21CNX.2 | 228 | -40 -86 | N
ULUF P 610 GG | 230V 50/60Hz | 8.9 | Green Gas | GWP3,9 |  | SC21CNX.2 / SC21CNX.2 | 242 | -40 -86 | N
ULUF P 820 GG | 230V 50/60Hz | 10.3 | Green Gas | GWP3,9 |  | SC21CNX.2 / SC21CNX.2 | 270 | -40 -86 | N

## kWh to BTU
Omregning af kWh to BTU (British thermal unit, heat output)
|  | 1 kWh = 3412.14 BTU |  |  |  | 1 BTU = 0.000293 kWh
|  | Convertion kWh to BTU: |  |  |  | Convertion BTU to kWh:
|  | 10 | kWh |  |  | 10 | BTU
|  | 34121.4163 | BTU |  |  | 0.0029307 | kWh
|  | 3412.1416 | BTU til 1 kWH
Varmeafgivelse
| år | Unit | Ambient °C | Stable temp °C | watt/h | kWh | BTU | BTU/h
| Katalog 2020 | ULUF 15 |  |  | 270.8333 | 6.5 | 22178.9206 | 924.1217
| Katalog 2020 | ULUF 65 |  |  | 287.5 | 6.9 | 23543.7773 | 980.9907
| 2020 | ULUF 125 | +21,5°C | -80°C | 283.125 | 6.795 | 23185.5024 | 966.0626
|  | ULUF 120 |  |  |  |  | 0 | 0 |  |  | 23884.9914 | 995.208
[7 empty] | 0 | 0
| 2020 |  | +21,5°C | -70°C |  | 5.412 | 18466.5105 | 769.4379
| 2020 | ULTF 80 | +20°C | -80°C |  | 5.5 | 18766.779 | 781.9491
| 2020 |  | +20°C | -70°C |  | 3.413 | 11645.6394 | 485.235
| 2020 | ULTF 420 | +20°C | -80°C |  | 8.26 | 28184.2899 | 1174.3454 |  |  | 10,6 kWh/24h
| 2020 |  | +20°C | -70°C |  | 6.862 | 23414.1159 | 975.5882 |  |  | P500GG: 12,4 kWh/24h
| 2018 | P390 | - | - |  | 15 | 51182.1245 | 2132.5885 |  |  | P610GG: 13,5 kWh/24h
| 2018 | P500 | - | - |  | 15.7 | 53570.6236 | 2232.1093 |  |  | P820GG: 17,3
| 2018 | P610 | - | - |  | 17.2 | 58688.8361 | 2445.3682
| 2018 | P820 | - | - |  | 20.9 | 71313.7601 | 2971.4067
| 2017 | CRYO 170 |  |  |  | 31.7 | 108164.8898 | 4506.8704
| 2017 | CRYO 230 | +25°C | -150°C |  | 33.5 | 114306.7447 | 4762.781
| 2020 | LRE 1400 | - | - |  | 2.8 | 9553.9966 | 398.0832 | [6 empty] | kWh | BTU | BTU/h
| Katalog 2020 | P10 |  |  |  | 4 | 13648.5665 | 568.6903 |  |  |  |  | Katalog 2020 | P10 | 4 | 13648.5665 | 568.6903
| Katalog 2020 | P50 |  |  |  | 5.6 | 19107.9931 | 796.1664 |  |  |  |  | Katalog 2020 | P50 | 5.6 | 19107.9931 | 796.1664
| Katalog 2020 | P90 |  |  |  | 7 | 23884.9914 | 995.208 | pr døgn | pr time |  |  | Katalog 2020 | P90 | 6.1 | 20814.064 | 867.2527
| Katalog 2021 | P390GG |  |  |  | 10.6 | 36168.7013 | 1507.0292 | 46046.4 | 1918.6 |  |  | Katalog 2020 | ULUF 15 | 6.5 | 22178.9206 | 924.1217
| Katalog 2021 | P500GG |  |  |  | 12.4 | 42310.5563 | 1762.9398 | 53865.6 | 2244.4
| Katalog 2021 | P610GG |  |  |  | 13.5 | 46063.912 | 1919.3297 | 58644 | 2443.5
| Katalog 2021 | P820GG |  |  |  | 17.3 | 59030.0503 | 2459.5854 | 75151.2 | 3131.3
| Website 2021 | LFTE 515 |  |  |  | 6.4 | 21837.7065 | 909.9044 | 27801.6 | 1158.4
| 2021 | PR 500 |  |  |  | 3.6 | 12283.7099 | 511.8212 | 15638.4 | 651.6
| 2021 | PR 300 |  |  |  | 3.4 | 11601.2816 | 483.3867 | 14769.6 | 615.4
| 2021 | LFF 270 |  |  |  | 11.4 | 38898.4146 | 1620.7673 | 49521.6 | 2063.4
| 2025 | P10MV |  |  |  | 2.35 | 8018.5328 | 334.1055
| 2025 | P50MV |  |  |  | 3.48 | 11874.2529 | 494.7605
| 2025 | P90MV |  |  |  | 4.67 | 15934.7014 | 663.9459
| 2025 | P830MV |  |  |  | 3.54 | 12078.9814 | 503.2909
|  | LR 1400 |  |  |  | 6.5 | 22178.9206 | 924.1217

## Voltage stabiliser
Voltage stabiliser
Model | Power supply | Kwh/24h | amps | watt
Cryo 170 | 400 | 31.7 | 4.8 | 1929
Cryo 230 | 400 | 33.5 | 4.7 | 2021
ULUF P10 | 230 | 4 | 3 | 370
ULUF P90 | 230 | 6.1 | 4 | 440
ULUF P390 | 230 | 15 | 8.7 | 1182
ULUF 820 | 230 | 20.9 | 9.2 | 1502
MUF 40 | 230 | 5.2 | 1.6 | 318
LTF | 230 | 5.1 | 2 | 297
LTFE 40 | 230 | 2.5 | 1.9 | 187
Stabiliser | Amps | VA at 230V | UK13 | EURO | US | Indian
SVS02 | 2 |  | x | x | x
SVS04 | 4 |  | x | x | x | x
SVS08 | 8 |  | x | x | x (115V)
SVS15 | 15 |  | x
SVS16 | 16 |  |  | x

## ST enheder Bolarus
Modtaget i mail d. 05.05.2020 Af Andrzej Bolarus CC Bartosz
D = Danfoss / E = Embraco)
x- available compressors with universal power supply 50/60 Hz
Universal compressors for Arctiko
|  |  |  | R404a / R134a 230V 50/60 HZ | [5 empty] | R290 230V 50Hz |  | R290 230V 60Hz |  | R290 115V 60Hz |  | The amount of R290 [g] | Capillary length
|  |  |  | Standard compressor (brand/model) | Standard compressor (brand/model) | Danfoss | Embraco | Additional Notifications | Tecumseh | Standard compressor (brand/model) | Additional Notifications | Standard compressor (brand/model) | Additional Notifications | Standard compressor (brand/model) | Additional Notifications
| Laboratory units - white |  | LR 55 | [6 empty] | HTD45AA
|  |  | LR 100 | [6 empty] | TL4CN, (Rep. EMT2117U) |  | EMI50UER |  | TL4.0CNX.2, EMI50UER |  | 80 | ø0,8 - 3,1m
|  |  | LF 100 | [6 empty] | SC12CNX2 (Rep. NEK2134U ) |  | NLE8.8CNT |  | SC12CNX.2 (Rep. NEK2134U) |  | 70 | ø1 - 3,6m
|  | S - 300 S | LR 300 | TL5G (D / 50/60 Hz), TL4G (D / 50 Hz) | TL4G |  |  |  |  | EMT2117U ( Rep. TL4CN) |  | EMI50UER |  | TL4.0CNX.2, EMI50UER |  | 90 | ø1 - 3,3m
|  | SN - 300 S | LF 300 | NEK2134GK (E / 50 Hz), SC12CL (D / 50 Hz) | [5 empty] | SC12CNX2 (Rep. NEK2134U ) |  | NLE8.8CNT (Rep. SCE15CNLX) |  | SC12CNX.2 (Rep.NEK2134U) |  | 100 | ø1 - 2,9m
|  | S/SN - 300 S | LFF 270 | TL5G, TL4G / NEK2134GK, SC12CL | [5 empty] | TL4CN (Rep. EMT2117U) / SC12CNX2 (Rep. NEK2134U) |  | EMI50UER / NT2160U(V), (Rep. SCE15CNLX) |  | TL4.0CNX.2 / SC12CNX.2 (Rep. NEK2134U) |  | 90, 110 | ø1 - 3,5m/ ø1 - 3,7m
|  | S/S - 300 S | LR 270-2 | TL5G, TL4G / TL5G, TL4G | [5 empty] | TL4CN (Rep. EMT2117U) |  | EMI50UER |  | TL4.0CNX.2, EMI50UER |  | 90, 90 | ø1 - 3,7m
|  |  | LFFG 270 | TL5G, TL4G / NEK2134GK, SC12CL | [5 empty] | TL4CN (Rep. EMT2117U) / SC12CNX2 (Rep. NEK2134U) |  | EMI50UER / NT2160U(V), (Rep. SCE15CNLX) |  | TL4.0CNX.2, EMI50UER / SC12CNX.2 (Rep. NEK2134U)
|  | S - 500 S | LR 500 | EMT6152GK, FR8.5G | [5 empty] | EMT2117U (Rep. TL5CN, EMT6152U) |  | EMI70UER |  | TL4.0CNX.2, EMI50UER |  | 110 | ø0,8 - 2,4m
|  | SN - 500 S | LF 500 | NEK2168GK, SC18CL, SC18CLX2 | [5 empty] | NT2160U (Rep. SC18CNX2) |  | NT2160U(V) (Rep. SCE18CNLX) |  | NT2160UV (Rep. SCE18CNLX) |  | 110 | ø1,2 - 2,25m
|  | S - 711 S | LR 700 | EMT6152GK, FR8.5G | [5 empty] | EMT 2117U (Rep. TL5CN, EMT6152U) |  | EMI70UER |  | TL4.8CNX.2 (Rep. EMT6152U) |  | 110 | ø0,8 - 2,4m
|  | SN - 711 S | LF 700 | NEK2168GK, SC18CL, SC18CLX2 | [5 empty] | NT2160U (Rep. SC18CNX2) |  | NT2160U(V), (Rep. SCE18CNLX) |  | NT2160UV (Rep. SCE18CNLX) |  | 110 | ø1,2 - 2,25m
|  |  | LAF 700 | [12 empty] | 150 | ø1,2 - 3,6m
|  |  | LFFG 660 | [6 empty] | EMT2117U (Rep. TL5CN,EMT6152U) / SC12CNX2 (Rep. NEK2134U) |  | EMI50UER / NLE8.8CNT, SCE15CNLX |  | TL4.0CNX.2 (Rep. EMT6152U) / NEK2160U (Rep. NEK2134UV SC12CNX.2) |  |  | ø
|  | S/SN - 711 S | LFF 660 | EMT6152GK, NEK2116Z / NEK2134GK, SC10CLX | [5 empty] | EMT2117U (Rep. TL5CN,EMT6152U) / SC12CNX2 (Rep. NEK2134U) |  | EMI50UER / NT2160U(V) (Rep. SCE15CNLX) |  | TL4.0CNX.2 (Rep. EMT6152U) / NEK2160U (Rep. NEK2134UV SC12CNX.2) |  | 90, 110 | ø0,8 - 2,25m/ ø1 - 3,7m
|  | S/S - 711 S | LR 660-2 | EMT6152GK, FR8.5G / EMT6152GK, FR8.5G | [5 empty] | EMT2117U (Rep. TL5CN, EMT2117U) |  | EMI50UER |  | TL4.0CNX.2, EMI50UER
|  |  | LR 900 | FR11G | [5 empty] | EMT2117U (Rep. TL5CN, EMT2117U) |  | EMI50UER |  | TL4.8CNX.2, EMI50UER
LF 900
|  | S - 147 S | LR 1400 | NEK6210GK, NEK2134GK, SC10CLX, SC10CL | [5 empty] | NEK6210U (Rep. NEK2134U) |  | NLE8.8CNT |  | NT6210U (Rep. NEK2134U) |  | 130 | ø1,2 - 2,25m
|  | SN - 147 S | LF 1400 | NT2192GK, GS26CLX | [5 empty] | NT2210U |  | NT2210UV |  | NT2210UV |  | 200 | ø1,2 - 2,3m
| Pharmaceutical units - white |  | PRE 55 | [6 empty] | HTD45AA
|  |  | PR 100 | [6 empty] | TL4CN, (Rep. EMT2117U) |  | EMI50UER |  | TL4.0CNX.2, EMI50UER |  | 80 | ø0,8 - 3,1m
|  |  | PF 100 | [6 empty] | SC12CNX2 (Rep. NEK2134U ) |  | NLE8.8CNT |  | SC12CNX.2 (Rep. NEK2134U) |  |  | ø1 - 3,6m
|  | WS - 300 S | PR 300 | EMT6152GK, FR8.5G | [5 empty] | EMT2117U ( Rep. TL4CN) |  | EMI50UER |  | TL4.0CNX.2, EMI50UER |  | 90 | ø1 - 3,3m
|  |  | PF 300 | [6 empty] | SC12CNX2 (Rep. NEK2134U ) |  | NLE8.8CNT (Rep. SCE15CNLX) |  | SC12CNX.2 (Rep.NEK2134U) |  | 100 | ø1 - 2,9m
|  | WS - 500 S | PR 500 | EMT6152GK, FR8.5G | [5 empty] | EMT2117U (Rep. TL5CN, EMT6152U) |  | EMI70UER |  | EMI70UER (Rep. EMT6152U, TL4.0CNX.2) |  | 110 | ø0,8 - 2,4m
|  | WSN - 500 S | PF 500 | NEK2168GK, SC18CL, SC18CLX2 | [5 empty] | NT2160U (Rep. SC18CNX2) |  | NT2160U(V) (Rep. SCE18CNLX) |  | NT2160UV (Rep. SCE18CNLX) |  | 110 | ø1,2 - 2,25m
|  | WS - 711 S | PR 700 | EMT6152GK, FR8.5G | [5 empty] | EMT 2117U (Rep. TL5CN, EMT6152U) |  | EMI70UER |  | TL4.8CNX.2 (Rep. EMT6152U) |  | 110 | ø0,8 - 2,4m
|  | WSN - 711 S | PF 700 | NEK2168GK, SC18CL, SC18CLX2 | [5 empty] | NT2160U (Rep. SC18CNX2) |  | NT2160U(V), (Rep. SCE18CNLX) |  | NT2160UV (Rep. SCE18CNLX) |  | 110 | ø1,2 - 2,25m
|  | WS - 147 S | PR 1400 | NEK6210GK, NEK2134GK, SC10CLX, SC10CL | [5 empty] | NEK6210U (Rep. NEK2134U) |  | NLE8.8CNT |  | NT6210U (Rep. NEK2134U) |  | 130 | ø1,2 - 2,25m
|  | WSN - 147 S | PF 1400 | NT2192GK, GS26CLX | [5 empty] | NT2210U |  | NT2210UV |  | NT2210UV |  | 200 | ø1,2 - 2,3m
| BBR Series |  | BBR 100 | [6 empty] | TL4CN (Rep. EMT2117U) |  | EMI50UER |  | TL4.0CNX.2, EMI50UER |  | 80 | ø0,8 - 3,1m
|  |  | BBR 300 | [6 empty] | EMT 2117U ( Rep. TL4CN) |  | EMI50UER |  | TL4.0CNX.2, EMI50UER |  | 90 | ø1 - 3,3m
|  |  | BBR 500 | [6 empty] | EMT 2117U (Rep. TL5CN, EMT6152U) |  | EMI70UER |  | TL4.0CNX.2, EMI50UER |  | 110 | ø0,8 - 2,4m
|  |  | BBR 700 | [6 empty] | EMT 2117U (Rep. TL5CN, EMT6152U) |  | EMI70UER |  | TL4.8CNX.2 (Rep. EMT6152U) |  | 110 | ø0,8 - 2,4m
|  |  | BBR 300 D | [6 empty] | EMT2117U (Rep. TL5CN) |  | EMI50UER |  | TL4.0CNX.2, EMI50UER |  | 90 | ø1 - 3,3m
|  |  | BBR 1400 | [6 empty] | NEK6210U (Rep. NEK2134U) |  | NLE8.8CNT |  | NT6210U (Rep. NEK2134U) |  | 130 | ø1,2 - 2,25m
|  |  | BBR 700 D | [6 empty] | EMT2117U, (Rep. EMI70UER) (-20/45) |  | EMI70UER |  | TL4.8CNX.2 |  | 90, 90 | ø0,8 - 2,4m/ ø0,8 - 2,4m
|  |  | BBR 1400 D | [6 empty] | NEK2134U (Rep. NEK6210U ) (-20/45) |  | NLE8.8CNT |  | NEK2134U, (Rep. NT6210U) |  | 80, 80 | ø1,2 - 2,25m

## GWP values
GPW values
Type | Refrigerant | Chemical name |  | GPW values
PFC | R-14 | Tetrafluoromethane |  | 7390
HFC | R-23 | Trifluoromethane (Fluoroform) |  | 14800 | [11 empty] | 3 | 3.3 | 12200 | 14800 | 7390
HC | R-50 | Methane |  | 25 | [11 empty] | R600a | R290 | R116 | R23 | R14
PFC | R-116 | Hexafluoroethane |  | 12200 | [11 empty] | 34 | 23 | 26 | 17 |  | 100 | 45
HFC | R-134a | 1,1,1,2-Tetrafluoroethane |  | 1430
HC | R-170 | Ethane |  | 5.5 | [11 empty] | 91.8 | 62.1 | 70.2 | 45.9 | 45 | 270
HC | R-290 | Propane |  | 3.3
HFC | R-404A | R-125/143a/134a (44±2/52±1/4±2) |  | 3922
HFC | R-508B | R-23/116 (46/54) |  | 13396
HC | R-600 | Butane |  | 4
HC | R-600a | Isobutane |  | 3
HO | R-1150 | Ethene (Ethylene) |  | 3.7 |  |  |  | GWP
HFO | R-1234yf | 2,3,3,3-Tetrafluoropropene |  | 4 |  |  | GWP Values | 3 | 3.3 | 12200 | 14800 | 7390 | [5 empty] | CO2 eq gasser udregnet i tons
HO | R-1270 | Propene (Propylene) |  | 1.8 |  |  |  | R600a | R290 | R116 | R23 | R14 | Total | GWP |  |  |  | R600a | R290 | R116 | R23 | R14
[7 empty] | % fyldning | 0.34 | 0.23 | 0.26 | 0.17 | 0.16667 | 1.1667 | EP88 | R14 | Total |  | 0.34 | 0.23 | 0.26 | 0.17 | 0.16667 | Total
MUF 40 | -86 | DTF 35 | EP88 | 138 | 23 | 161 |  | 46.92 | 31.74 | 35.88 | 23.46 | 23 | 161 | 5689.779 | 1231.6667 | 6921.4457 |  | 4.692e-05 | 3.174e-05 | 3.588e-05 | 2.346e-05 | 2.3e-05 | 0.000161
ULUF 450 | -86 | ULUF 450 | EP88 | 270 | 45 | 315 | 5945 | 91.8 | 62.1 | 70.2 | 45.9 | 45 | 315 | 5689.779 | 1231.6667 | 6921.4457 |  | 9.18e-05 | 6.21e-05 | 7.02e-05 | 4.59e-05 | 4.5e-05 | 0.000315
ULUF 550 | -86 | ULUF 550 | EP 88 | 290 | 68 | 358 |  | 98.6 | 66.7 | 75.4 | 49.3 | 68 | 358 | 5689.779 | 1231.6667 | 6921.4457 |  | 9.86e-05 | 6.67e-05 | 7.54e-05 | 4.93e-05 | 6.8e-05 | 0.000358
ULUF 750 | -86 | ULUF 750 | EP88 | 540 | (2*45) | 630 |  | 183.6 | 124.2 | 140.4 | 91.8 | 90 | 630 | 5689.779 | 1231.6667 | 6921.4457 |  | 0.0001836 | 0.0001242 | 0.0001404 | 9.18e-05 | 9e-05 | 0.00063
ULUF 15 | -86 | ULUF 15 | EP88 | 150 | 23 | 173 |  | 51 | 34.5 | 39 | 25.5 | 23 | 173 | 5689.779 | 1231.6667 | 6921.4457 |  | 5.1e-05 | 3.45e-05 | 3.9e-05 | 2.55e-05 | 2.3e-05 | 0.000173
ULUF 65 | -86 | ULUF 65 | EP88 | 171 | 24 | 195 |  | 58.14 | 39.33 | 44.46 | 29.07 | 24 | 195 | 5689.779 | 1231.6667 | 6921.4457 |  | 5.814e-05 | 3.933e-05 | 4.446e-05 | 2.907e-05 | 2.4e-05 | 0.000195
ULUF 125 | -86 | ULUF 125 | EP88 | 171 | 24 | 195 |  | 58.14 | 39.33 | 44.46 | 29.07 | 24 | 195 | 5689.779 | 1231.6667 | 6921.4457 |  | 5.814e-05 | 3.933e-05 | 4.446e-05 | 2.907e-05 | 2.4e-05 | 0.000195
[7 empty] | GWP værdi EP88 | 5689.779
[7 empty] | % af EP88 vs R14 | 0.85714 | 0.14286 |  | Samlet GWP værdi
[7 empty] | GWP fyldning af % | 4876.9534 | 1055.7143 |  | 5932.6677
Co2 eq
[7 empty] | ULUF 450 | 9.2e-05 | 6.2e-05 | 7e-05 | 4.6e-05 | 4.5e-05
[7 empty] | ULUF 450 | 0.000276 | 0.0002046 | 0.854 | 0.6808 | 0.33255 | 1.8678
ex. På typeskilt fra SHG d. 30/06-2020

## List of refrigerants
List of refrigerants
Chemical refrigerants are assigned an R number which is determined systematically according to molecular structure. Common refrigerants are frequently referred to as Freon (a registered trademark of DuPont). The following is a list of refrigerants with their Type/Prefix, ASHRAE designated numbers, IUPAC chemical name, molecular formula, CAS registry number / Blend Name, Atmospheric Lifetime in years, Semi-Empirical Ozone depletion potential, net Global warming potential over a 100-year time horizon, Occupational exposure limit/Permissible exposure limit in parts per million (volume per volume) over a time-weighted average (TWA) concentration for a normal eight-hour work day and a 40-hour work week, ASHRAE 34 Safety Group in Toxicity & Flammability (in Air @ 60 °C & 101.3 kPa) classing, Refrigerant Concentration Limit / Immediately Dangerous to Life or Health in parts per million (volume per volume) & grams per cubic meter, Molecular mass in Atomic mass units, Normal Boiling Point (or Bubble & Dew Points for the Zeotrope(400)-series)(or Normal Boiling Point & Azeotropic Temperature for the Azeotrope(500)-series) at 101,325 Pa (1
atmosphere) in degrees Celsius, Critical Temperature in degrees Celsius and Critical Pressure (absloute) in kiloPascals.
Contents
1 List
2 Type and flammability
3 See also
4 References
5 Notes
6 External links
List
T y p e | ASHRAE Number | IUPAC Chemical Name |  | net GWP 100-yr | OEL/ ppm ( ASHR Saf Gro
PCC | R-10 | Carbon tetrachloride (Tetrachloromethane) |  | 1,400[1] | 100?
CFC | R-11 | Trichlorofluoromethane |  | 4,750[1] | 1,00 A1
CFC | R-12 | Dichlorodifluoromethane |  | 10,900[1] | 1,00 A1
H | R-12B1 | Bromochlorodifluoromethane |  | 1,890[1]
H | R-12B2 | Dibromodifluoromethane
CFC | R-13 | Chlorotrifluoromethane |  | 14,400[1] | 1,00 A1
H | R-13B1 | Bromotrifluoromethane |  | 7,140[1] | 1,00 A1
PFC | R-14 | Tetrafluoromethane |  | 7,390[1] | 1,00 A1
HCC | R-20 | Chloroform (Trichloromethane) |  | 31[1]
HCFC | R-21 | Dichlorofluoromethane |  | 151[2] | 100?
HCFC | R-22 | Chlorodifluoromethane |  | 1,810[1] | 1,00 A1
H | R-22B1 | Bromodifluoromethane |  | 404[2]
HFC | R-23 | Trifluoromethane (Fluoroform) |  | 14,800[1] | 1,00 A1
HCC | R-30 | Dichloromethane (Methylene chloride) |  | 8.7[1] | 100?
HCFC | R-31 | Chlorofluoromethane |  |  | 350?
HFC | R-32 | Difluoromethane |  | 675[1] | 1,00 A2L
HCC | R-40 | Chloromethane |  | 13[1] | 100?
HFC | R-41 | Fluoromethane |  | 92[1]
HC | R-50 | Methane |  | 25[1] | 1,00 A3
PCC | R-110 | Hexachloroethane
CFC | R-111 | Pentachlorofluoroethane
CFC | R-112 | 1,1,2,2-Tetrachloro-1,2-difluoroethane
CFC | R-112a | 1,1,1,2-Tetrachloro-2,2-difluoroethane
CFC | R-113 | 1,1,2-Trichloro-1,2,2-trifluoroethane |  | 6,130[1] | 1,00 A1
CFC | R-113a | 1,1,1-Trichloro-2,2,2-trifluoroethane
CFC | R-114 | 1,2-Dichlorotetrafluoroethane |  | 10,000[1] | 1,00 A1
CFC | R-114a | 1,1-Dichlorotetrafluoroethane
H | R-114B2 | Dibromotetrafluoroethane |  | 1,640[1]
CFC | R-115 | Chloropentafluoroethane |  | 7,370[1] | 1,00 A1
PFC | R-116 | Hexafluoroethane |  | 12,200[1] | 1,00 A1
HCC | R-120 | Pentachloroethane
HCFC | R-121 | 1,1,2,2-Tetrachloro-1-fluoroethane
HCFC | R-121a | 1,1,1,2-Tetrachloro-2-fluoroethane
HCFC | R-122 | 1,1,2-Trichloro-2,2-difluoroethane
HCFC | R-122a | 1,1,2-Trichloro-1,2-difluoroethane
HCFC | R-122b | 1,1,1-Trichloro-2,2-difluoroethane
HCFC | R-123 | 2,2-Dichloro-1,1,1-trifluoroethane |  | 77[1] | 50[7]
HCFC | R-123a | 1,2-Dichloro-1,1,2-trifluoroethane |  | 77?[c]
HCFC | R-123b | 1,1-Dichloro-1,2,2-trifluoroethane |  | 77?[c]
HCFC | R-124 | 1-Chloro-1,2,2,2-tetrafluoroethane |  | 609[1] | 1,00 A1
HCFC | R-124a | 1-Chloro-1,1,2,2-tetrafluoroethane |  | 609?[c]
HFC | R-125 | Pentafluoroethane |  | 3,500[1] | 1,00 A1
HFC | R-E125 | Pentafluorodimethyl ether |  | 14,900[1]
HCC | R-130 | 1,1,2,2-Tetrachloroethane
HCC | R-130a | 1,1,1,2-Tetrachloroethane
HCFC | R-131 | 1,1,2-Trichloro-2-fluoroethane
HCFC | R-131a | 1,1,2-Trichloro-1-fluoroethane
HCFC | R-131b | 1,1,1-Trichloro-2-fluoroethane
HCFC | R-132 | Dichlorodifluoroethane
HCFC | R-132a | 1,1-Dichloro-2,2-difluoroethane
HCFC | R-132b | 1,2-Dichloro-1,1-difluoroethane
HCFC | R-132c | 1,1-Dichloro-1,2-difluoroethane
H | R-132bB2 | 1,2-Dibromo-1,1-difluoroethane
HCFC | R-133 | 1-Chloro-1,2,2-Trifluoroethane
HCFC | R-133a | 1-Chloro-2,2,2-Trifluoroethane
HCFC | R-133b | 1-Chloro-1,1,2-Trifluoroethane
HFC | R-134 | 1,1,2,2-Tetrafluoroethane |  | 1,100[1]
HFC | R-134a | 1,1,1,2-Tetrafluoroethane |  | 1,430[1] | 1,00 A1
HFC | R-E134 | Bis(difluoromethyl)ether |  | 6,320[1]
HCC | R-140 | 1,1,2-Trichloroethane
HCC | R-140a | 1,1,1-Trichloroethane (Methyl chloroform) |  | 146[1]
HCFC | R-141 | 1,2-Dichloro-1-fluoroethane |  | 725?[c]
H | R-141B2 | 1,2-Dibromo-1-fluoroethane
HCFC | R-141a | 1,1-Dichloro-2-fluoroethane |  | 725?[c]
HCFC | R-141b | 1,1-Dichloro-1-fluoroethane |  | 725[1] | 500[7]
HCFC | R-142 | Chlorodifluoroethane |  | 2,310?[c]
HCFC | R-142a | 1-Chloro-1,2-difluoroethane |  | 2,310?[c]
HCFC | R-142b | 1-Chloro-1,1-difluoroethane |  | 2,310[1] | 1,00 A2
HFC | R-143 | 1,1,2-Trifluoroethane |  | 353[1]
HFC | R-143a | 1,1,1-Trifluoroethane |  | 4,470[1] | 1,00 A2L
HFC | R-143m | Methyl trifluoromethyl ether
HFC | R-E143a | 2,2,2-Trifluoroethyl methyl ether |  | 756[1]
HCC | R-150 | 1,2-Dichloroethane
HCC | R-150a | 1,1-Dichloroethane
HCFC | R-151 | Chlorofluoroethane
HCFC | R-151a | 1-Chloro-1-fluoroethane
HFC | R-152 | 1,2-Difluoroethane |  | 53[1]
HFC | R-152a | 1,1-Difluoroethane |  | 124[1] | 1,00 A2
HCC | R-160 | Chloroethane (ethyl chloride)
HFC | R-161 | Fluoroethane |  | 12[1]
HC | R-170 | Ethane |  | 5.5[1] | 1,00 A3
HC | R-E170 | Dimethyl ether |  | 1[1] | 1,00 A3
CFC | R-211 | 1,1,1,2,2,3,3-Heptachloro-3-fluoropropane
CFC | R-212 | Hexachlorodifluoropropane
CFC | R-213 | 1,1,1,3,3-Pentachloro-2,2,3-trifluoropropane
CFC | R-214 | 1,2,2,3-Tetrachloro-1,1,3,3-tetrafluoropropane
CFC | R-215 | 1,1,1-Trichloro-2,2,3,3,3-pentafluoropropane
CFC | R-216 | 1,2-Dichloro-1,1,2,3,3,3-hexafluoropropane
CFC | R-216ca | 1,3-Dichloro-1,1,2,2,3,3-hexafluoropropane
CFC | R-217 | 1-Chloro-1,1,2,2,3,3,3-heptafluoropropane
CFC | R-217ba | 2-Chloro-1,1,1,2,3,3,3-heptafluoropropane
PFC | R-218 | Octafluoropropane |  | 8,830[1] | 1,00 A1
HCFC | R-221 | 1,1,1,2,2,3-Hexachloro-3-fluoropropane
HCFC | R-222 | Pentachlorodifluoropropane
HCFC | R-222c | 1,1,1,3,3-Pentachloro-2,2-difluoropropane
HCFC | R-223 | Tetrachlorotrifluoropropane
HCFC | R-223ca | 1,1,3,3-Tetrachloro-1,2,2-trifluoropropane
HCFC | R-223cb | 1,1,1,3-Tetrachloro-2,2,3-trifluoropropane
HCFC | R-224 | Trichlorotetrafluoropropane
HCFC | R-224ca | 1,3,3-Trichloro-1,1,2,2-tetrafluoropropane
HCFC | R-224cb | 1,1,3-Trichloro-1,2,2,3-tetrafluoropropane
HCFC | R-224cc | 1,1,1-Trichloro-2,2,3,3-tetrafluoropropane
HCFC | R-225 | Dichloropentafluoropropane
HCFC | R-225aa | 2,2-Dichloro-1,1,1,3,3-pentafluoropropane
HCFC | R-225ba | 2,3-Dichloro-1,1,1,2,3-pentafluoropropane
HCFC | R-225bb | 1,2-Dichloro-1,1,2,3,3-pentafluoropropane
HCFC | R-225ca | 3,3-Dichloro-1,1,1,2,2-pentafluoropropane |  | 122[1]
HCFC | R-225cb | 1,3-Dichloro-1,1,2,2,3-pentafluoropropane |  | 595[1]
HCFC | R-225cc | 1,1-Dichloro-1,2,2,3,3-pentafluoropropane
HCFC | R-225da | 1,2-Dichloro-1,1,3,3,3-pentafluoropropane
HCFC | R-225ea | 1,3-Dichloro-1,1,2,3,3-pentafluoropropane
HCFC | R-225eb | 1,1-Dichloro-1,2,3,3,3-pentafluoropropane
HCFC | R-226 | Chlorohexafluoropropane
HCFC | R-226ba | 2-Chloro-1,1,1,2,3,3-hexafluoropropane
HCFC | R-226ca | 3-Chloro-1,1,1,2,2,3-hexafluoropropane
HCFC | R-226cb | 1-Chloro-1,1,2,2,3,3-hexafluoropropane
HCFC | R-226da | 2-Chloro-1,1,1,3,3,3-hexafluoropropane
HCFC | R-226ea | 1-Chloro-1,1,2,3,3,3-hexafluoropropane
HFC | R-227ca | 1,1,2,2,3,3,3-Heptafluoropropane
HFC | R-227ca2 | Trifluoromethyl 1,1,2,2-tetrafluoroethyl ether
HFC | R-227ea | 1,1,1,2,3,3,3-Heptafluoropropane |  | 3,220[1] | 1,00 A1
HFC | R-227me | Trifluoromethyl 1,2,2,2-tetrafluoroethyl ether |  | 1,540[2]
HCFC | R-231 | Pentachlorofluoropropane
HCFC | R-232 | Tetrachlorodifluoropropane
HCFC | R-232ca | 1,1,3,3-Tetrachloro-2,2-difluoropropane
HCFC | R-232cb | 1,1,1,3-Tetrachloro-2,2-difluoropropane
HCFC | R-233 | Trichlorotrifluoropropane
HCFC | R-233ca | 1,1,3-Trichloro-2,2,3-trifluoropropane
HCFC | R-233cb | 1,1,3-Trichloro-1,2,2-trifluoropropane
HCFC | R-233cc | 1,1,1-Trichloro-2,2,3-trifluoropropane
HCFC | R-234 | Dichlorotetrafluoropropane
HCFC | R-234aa | 2,2-Dichloro-1,1,3,3-tetrafluoropropane
HCFC | R-234ab | 2,2-Dichloro-1,1,1,3-tetrafluoropropane
HCFC | R-234ba | 1,2-Dichloro-1,2,3,3-tetrafluoropropane
HCFC | R-234bb | 2,3-Dichloro-1,1,1,2-tetrafluoropropane
HCFC | R-234bc | 1,2-Dichloro-1,1,2,3-tetrafluoropropane
HCFC | R-234ca | 1,3-Dichloro-1,2,2,3-tetrafluoropropane
HCFC | R-234cb | 1,1-Dichloro-2,2,3,3-tetrafluoropropane
HCFC | R-234cc | 1,3-Dichloro-1,1,2,2-tetrafluoropropane
HCFC | R-234cd | 1,1-Dichloro-1,2,2,3-tetrafluoropropane
HCFC | R-234da | 2,3-Dichloro-1,1,1,3-tetrafluoropropane
HCFC | R-234fa | 1,3-Dichloro-1,1,3,3-tetrafluoropropane
HCFC | R-234fb | 1,1-Dichloro-1,3,3,3-tetrafluoropropane
HCFC | R-235 | Chloropentafluoropropane
HCFC | R-235ca | 1-Chloro-1,2,2,3,3-pentafluoropropane
HCFC | R-235cb | 3-Chloro-1,1,1,2,3-pentafluoropropane
HCFC | R-235cc | 1-Chloro-1,1,2,2,3-pentafluoropropane
HCFC | R-235da | 2-Chloro-1,1,1,3,3-pentafluoropropane
HCFC | R-235fa | 1-Chloro-1,1,3,3,3-pentafluoropropane
HFC | R-236cb | 1,1,1,2,2,3-Hexafluoropropane |  | 1,340[1]
HFC | R-236ea | 1,1,1,2,3,3-Hexafluoropropane |  | 1,370[1]
HFC | R-236fa | 1,1,1,3,3,3-Hexafluoropropane |  | 9,810[1] | 1,00 A1
HFC | R-236me | 1,2,2,2-Tetrafluoroethyl difluoromethyl ether |  | 989[2]
HFC | R-FE-36 | Hexafluoropropane
HCFC | R-241 | Tetrachlorofluoropropane
HCFC | R-242 | Trichlorodifluoropropane
HCFC | R-243 | Dichlorotrifluoropropane
HCFC | R-243ca | 1,3-Dichloro-1,2,2-trifluoropropane
HCFC | R-243cb | 1,1-Dichloro-2,2,3-trifluoropropane
HCFC | R-243cc | 1,1-Dichloro-1,2,2-trifluoropropane
HCFC | R-243da | 2,3-Dichloro-1,1,1-trifluoropropane
HCFC | R-243ea | 1,3-Dichloro-1,2,3-trifluoropropane
HCFC | R-243ec | 1,3-Dichloro-1,1,2-trifluoropropane
HCFC | R-244 | Chlorotetrafluoropropane
HCFC | R-244ba | 2-Chloro-1,2,3,3-tetrafluoropropane
HCFC | R-244bb | 2-Chloro-1,1,1,2-tetrafluoropropane
HCFC | R-244ca | 3-Chloro-1,1,2,2-tetrafluoropropane
HCFC | R-244cb | 1-Chloro-1,2,2,3-tetrafluoropropane
HCFC | R-244cc | 1-Chloro-1,1,2,2-tetrafluoropropane
HCFC | R-244da | 2-Chloro-1,1,3,3-tetrafluoropropane
HCFC | R-244db | 2-Chloro-1,1,1,3-tetrafluoropropane
HCFC | R-244ea | 3-Chloro-1,1,2,3-tetrafluoropropane
HCFC | R-244eb | 3-Chloro-1,1,1,2-tetrafluoropropane
HCFC | R-244ec | 1-Chloro-1,1,2,3-tetrafluoropropane
HCFC | R-244fa | 3-Chloro-1,1,1,3-tetrafluoropropane
HCFC | R-244fb | 1-Chloro-1,1,3,3-tetrafluoropropane
HFC | R-245ca | 1,1,2,2,3-Pentafluoropropane |  | 693[1]
HFC | R-245cb | Pentafluoropropane
HFC | R-245ea | 1,1,2,3,3-Pentafluoropropane
HFC | R-245eb | 1,1,1,2,3-Pentafluoropropane
HFC | R-245fa | 1,1,1,3,3-Pentafluoropropane |  | 1,030[1] | 300[7]
HFC | R-245mc | Methyl pentafluoroethyl ether
HFC | R-245mf | Difluoromethyl 2,2,2-trifluoroethyl ether |  | 286[2]
HFC | R-245qc | Difluoromethyl 1,1,2-trifluoroethyl ether
HCFC | R-251 | Trichlorofluoropropane
HCFC | R-252 | Dichlorodifluoropropane
HCFC | R-252ca | 1,3-Dichloro-2,2-difluoropropane
HCFC | R-252cb | 1,1-Dichloro-2,2-difluoropropane
HCFC | R-252dc | 1,2-Dichloro-1,1-difluoropropane
HCFC | R-252ec | 1,1-Dichloro-1,2-difluoropropane
HCFC | R-253 | Chlorotrifluoropropane
HCFC | R-253ba | 2-Chloro-1,2,3-trifluoropropane
HCFC | R-253bb | 2-Chloro-1,1,2-trifluoropropane
HCFC | R-253ca | 1-Chloro-2,2,3-trifluoropropane
HCFC | R-253cb | 1-Chloro-1,2,2-trifluoropropane
HCFC | R-253ea | 3-Chloro-1,1,2-trifluoropropane
HCFC | R-253eb | 1-Chloro-1,2,3-trifluoropropane
HCFC | R-253ec | 1-Chloro-1,1,2-trifluoropropane
HCFC | R-253fa | 3-Chloro-1,3,3-trifluoropropane
HCFC | R-253fb | 3-Chloro-1,1,1-trifluoropropane
HCFC | R-253fc | 1-Chloro-1,1,3-trifluoropropane
HFC | R-254cb | 1,1,2,2-Tetrafluoropropane
HFC | R-254pc | Methyl 1,1,2,2-tetrafluoroethyl ether |  | 359[1]
HCFC | R-261 | Dichlorofluoropropane
HCFC | R-261ba | 1,2-Dichloro-2-fluoropropane
HCFC | R-262 | Chlorodifluoropropane
HCFC | R-262ca | 1-Chloro-2,2-difluoropropane
HCFC | R-262fa | 3-Chloro-1,1-difluoropropane
HCFC | R-262fb | 1-Chloro-1,3-difluoropropane
HFC | R-263 | Trifluoropropane
HCFC | R-271 | Chlorofluoropropane
HCFC | R-271b | 2-Chloro-2-fluoropropane
HCFC | R-271d | 2-Chloro-1-fluoropropane
HCFC | R-271fb | 1-Chloro-1-fluoropropane
HFC | R-272 | Difluoropropane
HFC | R-281 | Fluoropropane
HC | R-290 | Propane |  | 3.3[1] | 1,00 A3
CFC | R-C316 | Dichlorohexafluorocyclobutane
CFC | R-C317 | Chloroheptafluorocyclobutane
PFC | R-C318 | Octafluorocyclobutane (Perfluorocyclobutane) |  | 10,300[1] | 1,00 A1
PFC | R-3-1-10 | Decafluorobutane (Perfluorobutane) |  | 8,860[1]
HFC | R-329ccb | 1,1,1,2,2,3,3,4,4-Nonafluorobutane
HFC | R-338eea | 1,1,1,2,3,4,4,4-Octafluorobutane
HFC | R-347ccd | 1,1,1,2,2,3,3-Heptafluorobutane
HFC | R-347mcc | Perfluoropropyl methyl ether |  | 575[1]
HFC | R-347mmy | Perfluoroisopropyl methyl ether |  | 343[2]
HFC | R-365mfc | 1,1,1,3,3-Pentafluorobutane |  | 794[1]
PFC | R-4-1-12 | Dodecafluoropentane (Perfluoropentane) |  | 9,160[1]
PFC | R-5-1-14 | Tetradecafluorohexane (Perfluorohexane) |  | 9,300[1]
CFC | R-400 | R-12/114(50/50 wt%) OR (60/40)(must be specified) |  | 10,450[note 1] 10,540[note 1] | 1,00 A1 1,00 A1
HCFC | R-401A | R-22/152a/124 (53±2/13+.5,-1.5/34±1) |  | 1,182[note 1] | 1,00 A1
HCFC | R-401B | R-22/152a/124 (61±2/11+.5,-1.5/28±1) |  | 1,288[note 1] | 1,00 A1
HCFC | R-401C | R-22/152a/124 (33±2/15+.5,-1.5/52±1) |  | 933[note 1] | 1,00 A1
HCFC | R-402A | R-125/290/22 (60±2/2±1/38±2) |  | 2,788[note 1] | 1,00 A1
HCFC | R-402B | R-125/290/22 (38±2/2±1/60±2) |  | 2,416[note 1] | 1,00 A1
HCFC | R-403A | R-290/22/218 (5+.2,-2/75±2/20±0) |  | 3,124[note 1] | 1,00 A2[
HCFC | R-403B | R-290/22/218 (5+.2,-2/56±2/39±0) |  | 4,457[note 1] | 1,00 A1
HFC | R-404A | R-125/143a/134a (44±2/52±1/4±2) |  | 3,922[note 1] | 1,00 A1
HCFC | R-405A | R-22/152a/142b/C318 (45±0/7±1/5.5±1/42.5±2) |  | 5,328[note 1] | 1,000[
HCFC | R-406A | R-22/600a/142b (55±2/4±1/41±0) |  | 1,943[note 1] | 1,00 A2
HCFC | R-406B | R-22/600a/142b (65±2/4±1/31±0) |  | 1,893[note 1] | 1000
HFC | R-407A | R-32/125/134a (20±2/40±2/40±2) |  | 2,107[note 1] | 1,00 A1
HFC | R-407B | R-32/125/134a (10±2/70±2/20±2) |  | 2,804[note 1] | 1,00
A1
HFC | R-407C | R-32/125/134a (23±2/25±2/52±2) |  | 1,774[note 1]
1,00 A1
HFC | R-407D | R-32/125/134a (15±2/15±2/70±2) |  | 1,627[note 1] | 1,00 A1
HFC | R-407E | R-32/125/134a (25±2/15±2/60±2) |  | 1,552[note 1] | 1,00 A1
HFC | R-407F | R-32/125/134a (30±2/30±2/40±2) |  | 1,825[note 1] | 1,00 A1[
HCFC | R-408A | R-125/143a/22 (7±2/46±1/47±2) |  | 3,152[note 1] | 1,00 A1
HCFC | R-409A | R-22/124/142b (60±2/25±2/15±1) |  | 1,585[note 1] | 1,00 A1
HCFC | R-409B | R-22/124/142b (65±2/25±2/10±1) |  | 1,560[note 1] | 1,00 A1
HFC | R-410A | R-32/125 (50+.5,–1.5/50+1.5,–.5) |  | 2,088[note 1] | 1,00 A1
HFC | R-410B | R-32/125 (45±1/55±1) |  | 2,229[note 1] | 1,000?
HCFO | R-411A | R-1270/22/152a (1.5+0,–1/87.5+2,–0/11+0,–1) |  | 1,597[note 1] | 990[7]
HCFO | R-411B | R-1270/22/152a (3+0,–1/94+2,–0/3+0,–1) |  | 1,705[note 1] | 980[7]
HCFO | R-411C | R-1270/22/152a (3+0,–1/95.5+2,–0/1.5+0,–1) |  | 1,730[note 1] | 970?
HCFC | R-412A | R-22/218/142b (70±2/5±2/25±1) |  | 2,286[note 1] | 1,00 A2
HFC | R-413A | R-218/134a/600a (9±1/88±2/3+0,–1) |  | 2,053[note 1] | 1,00 A2
HCFC | R-414A | R-22/124/600a/142b (51±2/28.5±2/4±.5/16.5+.5,–1) |  | 1,478[note 1] | 1,00 A1
HCFC | R-414B | R-22/124/600a/142b (50±2/39±2/1.5±.5/9.5+.5,– 1) |  | 1,362[note 1] | 1,00 A1
HCFC | R-415A | R-22/152a (82±1/18±1) |  | 1,507[note 1] | 1,00 A2
HCFC | R-415B | R-22/152a (25±1/75±1) |  | 546[note 1] | 1,00 A2
HCFC | R-416A | R-134a/124/600 (59+.5,– 1/39.5+1,–.5/1.5+1,–.2) |  | 1,084[note 1] | 1,00 A1
HFC | R-417A | R-125/134a/600 (46.6±1.1/50±1/3.4+.1,–.4) |  | 2,346[note 1] | 1,00 A1
HFC | R-417B | R-125/134a/600 (79±1/18.3±1/2.7+.1,–.5) |  | 3,027[note 1] | 1,00 A1[
HCFC | R-418A | R-290/22/152a (1.5±.5/96±1/2.5±.5) |  | 1,741[note 1] | 1,00 A2
HFC | R-419A | R-125/134a/E170 (77±1/19±1/4±1) |  | 2,967[note 1] | 1,00 A2
HCFC | R-420A | R-134a/142b (88+1,–0/12+0,–1) |  | 1,548[note 1] | 1,00 A1
HFC | R-421A | R-125/134a (58±1/42±1) |  | 2,631[note 1] | 1,00 A1
HFC | R-421B | R-125/134a (85±1/15±1) |  | 3,190[note 1] | 1,00 A1
HFC | R-422A | R-125/134a/600a (85.1±1/11.5±1/3.4+.1,–.4) |  | 3,143[note 1] | 1,00 A1
HFC | R-422B | R-125/134a/600a (55±1/42±1/3+.1,–.5) |  | 2,526[note 1] | 1,00 A1
HFC | R-422C | R-125/134a/600a (82±1/15±1/3+.1,–.5) |  | 3,085[note 1] | 1,00 A1
HFC | R-422D | R-125/134a/600a (65.1+.9,– 1.1/31.5±1/3.4+.1,–.4) |  | 2,729[note 1] | 1,00 A1
HFC | R-423A | R-134a/227ea (52.5±1/47.5±1) |  | 2,280[note 1] | 1,00 A1
HFC | R-424A | R-125/134a/600a/600/601a (50.5±1/47±1/.9+.1,–.2/1+.1,+.2/.6+.1,–.2) |  | 2,440?[note 1] | 970[7]
HFC | R-425A | R-32/134a/227ea (18.5±.5/69.5±.5/12±.5) |  | 1,505[note 1] | 1,00 A1
HFC | R-426A | R-125/134a/600/601a (5.1±1/93±1/1.3+.1,–.2/.6+.1,–.2) |  | 1,508?[note 1] | 990[7]
HFC | R-427A | R-32/125/143a/134a (15±2/25±2/10±2/50±2) |  | 2,138[note 1] | 1,00 A1
HFC | R-428A | R-125/143a/290/600a (77.5±1/20±1/.6+.1,–.2/1.9+.1,–.2) |  | 3,607[note 1] | 1,00 A1
HFC | R-429A | R-E170/152a/600a (60±1/10±1/30±1) |  | 13.9[note 1] | 1,00 A3
HFC | R-430A | R-152a/600a (76±1/24±1) |  | 95[note 1] | 1,00 A3
HFC | R-431A | R-290/152a (71±1/29±1) |  | 38.3[note 1] | 1,00 A3
HO | R-432A | R-1270/E170 (80±1/20±1) |  | 1.64[note 1] | 710[7]
HO | R-433A | R-1270/290 (30±1/70±1) |  | 2.85[note 1] | 880[7]
HO | R-433B | R-1270/290 (5±1/95±1) |  | 3.23[note 1] | 950 A3[
HO | R-433C | R-1270/290 (25±1/75±1) |  | 2.93[note 1] | 790 A3[
HFC | R-434A | R-125/143a/134a/600a (63.2±1/18±1/16±1/2.8+.1,–.2) |  | 3,245[note 1] | 1,00 A1
HFC | R-435A | R-E170/152a (80±1/20±1) |  | 25.6[note 1] | 1,00 A3
HC | R-436A | R-290/600a (56±1/44±1) |  | 3.17[note 1] | 1,00 A3
HC | R-436B | R-290/600a (52±1/48±1) |  | 3.16[note 1] | 1,00 A3
HFC | R-437A | R-125/134a/600/601 (19.5+.5,– 1.8/78.5+1.5,–.7/1.4+.1,–.2/.6+.1,–.2) |  | 1,805?[note 1] | 990[7]
HFC | R-438A | R-32/125/134a/600/601a (8.5+.5,-1.5/45±1.5/44.2±1.5/1.7+.1,-.2/.6+.1,-.2) |  | 2,265?[note 1] | 990 A1[
HFC | R-439A | R-32/125/600a (50±1/47±1/3±.5) |  | 1,983[note 1] | 990 A2[
HFC | R-440A | R-290/134a/152a (.6±.1/1.6±.6/97.8±.5) |  | 144[note 1] | 1,00 A2[
HC | R-441A | R-170/290/600a/600 (3.1±.3/54.8±2/6±.6/36.1±2) |  | 3.6[note 1] | 1,00 A3[
HCFC | R-500 | R-12/152a (73.8/26.2) |  | 8,077[note 1] | 1,00 A1
HCFC | R-501 | R-22/12 (75/25) |  | 4,083[note 1] | 1,00 A1
CFC | R-502 | R-22/115 (48.8/51.2) |  | 4,657[note 1] | 1,00 A1
HCFC | R-503 | R-23/13 (40.1/59.9) |  | 14,560[note 1] | 1,000[
HCFC | R-504 | R-32/115 (48.2/51.8) |  | 4,143[note 1] | 1,000[
HCFC | R-505 | R-12/31 (78/22) |  | 8,504+[note 1] | 350?
HCFC | R-506 | R-31/114 (55.1/44.9) |  | 4,495+[note 1] | 350?
HFC | R-507[A] | R-125/143a (50/50) |  | 3,985[note 1] | 1,00 A1
HFC | R-508[A] | R-23/116 (39/61) |  | 13,214[note 1] | 1,00 A1
HFC | R-508B | R-23/116 (46/54) |  | 13,396[note 1] | 1,00 A1
HCFC | R-509[A] | R-22/218 (44/56) |  | 5,741[note 1] | 1,00 A1
HC | R-510[A] | R-E170/600a (88±.5/12±.5) |  | 1.24[note 1] | 1,00 A3
HC | R-511[A] | R-290/E170 (95±1/5±1) |  | 3.19[note 1] | 1,00 A3[
HC | R-600 | Butane |  | 4.0[1] | 1,00 A3
HC | R-600a | Isobutane |  | 3[28] | 1,00 A3
HC | R-601 | Pentane |  | 4 ± 2[f] | 600[7]
HC | R-601a | Isopentane |  | 4 ± 2[f] | 600[7]
HC | R-610 | Ethoxyethane (Diethyl ether) |  | 4 ± 2[f] | 400[7]
HC | R-611 | Methyl formate |  | < 25[11][28] | 100[7]
| R-630 | Methylamine
| R-631 | Ethylamine
| R-702 | Hydrogen |  | 5.8[1] | A3
| R-704 | Helium |  | 0[g] | 9,340?
| R-717 | Ammonia |  | 0[29] | 25[7] B
| R-718 | Water/Steam |  | 0.2 ± 0.2[1] | A1
| R-720 | Neon |  | 0[g] | 9,340?
| R-728 | Nitrogen |  | 0[g] | A1
| R-729 | Air N2/O2/Ar (78.082/20.945/.934 vol%) |  | 0.00054+[1] | A1
| R-732 | Oxygen |  | 0[g] | A1
| R-740 | Argon |  | 0[g] | 9,340?
| R-744 | Carbon dioxide |  | 1[1] | 5,00 A1
| R-744A | Nitrous oxide |  | 298[1] | A1
| R-764 | Sulfur dioxide |  |  | 5[33]
| R-784 | Krypton |  | 0[g] | 9340
CFO | R-1112a | 1,1-Dichloro-2,2-difluoroethylene
CFO | R-1113 | Chlorotrifluoroethylene
PFO | R-1114 | Tetrafluoroethylene
HCO | R-1120 | Trichloroethylene (trielene)
HCO | R-1130 | cis-1,2-Dichloroethylene |  | < 25[11]
HFO | R-1132a | 1,1-Difluoroethylene
HCO | R-1140 | Chloroethylene (vinyl chloride) |  |  | 100?
HFO | R-1141 | Fluoroethylene (vinyl fluoride)
HO | R-1150 | Ethene (Ethylene) |  | 3.7[1] | 200[7]
PFO | R-1216 | Hexafluoropropylene |  | 17,340+[1]
PFO | R-1218 | Hexafluoropropene trimer
HCFO | R-1233zd | 1-Chloro-3,3,3-trifluoropropene |  | 1[34]
HFO | R-1234yf | 2,3,3,3-Tetrafluoropropene |  | 4[28][35] | 500 A2L
HFO | R-1234ze | 1,3,3,3-Tetrafluoropropene |  | 6[28] | 800 A2L
HO | R-1270 | Propene (Propylene) |  | 1.8[1] | 500[7]
OEL/
Type | ASHRAE Number | IUPAC Chemical Name |  | net GWP 100-yr | ppm ( ASHR Saf
Gro
Type and flammability
Legends
Types / Prefixes[38] Safety Classification of Refrigerants[39]
Type | Meaning |  | Atoms in the Molecule
CFC | Chlorofluorocarbon |  | Cl, F, C | B3
CFO | Chlorofluoroolefin
HCFC | Hydrochlorofluorocarbon |  | H, Cl, F, C
B2
HCFO | Hydrochlorofluoroolefin
HFC | Hydrofluorocarbon |  | H, F, C
B2L
HFO | Hydrofluoroolefin
HCC | Hydrochlorocarbon |  | H, Cl, C
HCO | Hydrochloroolefin
B1
HC | Hydrocarbon |  | H, C
Higher Toxicity OEL < 400 ppm[23]
HO | Hydroolefin (Alkene)
PFC | Perfluorocarbon |  | F, C
PFO | Perfluoroolefin
PCC | Perchlorocarbon |  | Cl, C
PCO | Perchloroolefin
H | Halon/Haloalkane |  | Br, Cl (in some but not all), F, H (in some but not all), C
Compounds used as refrigerants may be described using either the appropriate prefix above or with the prefixes "R-" or "Refrigerant." Thus, CFC-12 may also be written as R-12 or Refrigerant 12.
An alkene, olefin, or olefine is anunsaturated chemical compoundcontaining at least onecarbon-to-carbon double bond.[41]
Educated Estimates
a. Some CFCs are assumed to be similar in terms of Semi-Empirical ODP being one.[2][8]
b. Some HCFCs are assumed to be similar in terms of Semi-Empirical OD.P[2][10]
c. Some HCFCs are assumed to be similar in terms of Atmospheric Lifetim[1e][2] & net GWP over a 100-year time horizon[.1][2]
d. All HFCs, HFOs, PFCs & PFOs are assumed to be similar in terms of Semi-Empirical ODP being zero.[9][11]
e. Most hydrocarbons are assumed to be similar to Methane in terms of Atmospheric Lifetime being 12 ± 3 years[1] & Semi-Empirical ODP being < 0(smog chemistry)[1. 4]
f. Pentane, Isopentane & Ethoxyethane (Diethyl ether) is assumed to have a net GWP over 100-yr to be 4 ± 2.
g. Air, nitrogen, oxygen & thenoble gases are assumed to have an unlimited Atmospheric Lifetime, a Semi-Empirical ODP of zero (except oxygen being slightly less than zero) & net GWP over 100-yr to be zero (except air being the sum of its parts).
See also
Haloalkane
References
1. P. Forster; V. Ramaswamy; P. Artaxo; T. Berntsen; R. Betts; D.W. Fahey; J. Haywood;
J. Lean; D.C. Lowe; G. Myhre; J. Nganga; R. Prinn; G. Raga; M. Schulz; R.aVn Dorland (2007). "Chapter 2: Changes in Atmospheric Constituents and in Radiative Forcing". In Solomon, S.; Mille,r H.L.; Tignor, M.; Averyt, K.B.; Marquis, M.; Chen, Z.; Manning, M.; Qin, D.Climate Change 2007: The Physical Science Basis. Contribution of Working Group I to the Fourth Assessment Report of the Intergovernmental Panel on Climate Change(https://www.ipcc.ch/publications_and_data/ar4/wg1/en/ch2.html). Cambridge, United Kingdom and New York, NY, USA: Cambridge University Press. Retrieved 9 October 2016.
2. John S. Daniel; Guus J.M. Velders; A.R. Douglass; P.M.D. Forster; D.A. Hauglustaine;
I.S.A. Isaksen; L.J.M. Kuijpers; A. McCulloch; .TJ. Wallington (2006). "Chapter 8. Halocarbon Scenarios, Ozone Depletion Potentials, and Global Warming Potentials". Scientific Assessment of Ozone Depletion: 2006(http://www.esrl.noaa.gov/csd/asses sments/ozone/2006/chapters/chapter8.pdf)(pdf). Geneva, Switzerland:World Meteorological Organization. Retrieved 9 October 2016.
3. "Addenda to ANSI/ASHRAE Standard 34-1997"(https://ashrae.org/standards-researc h--technology/standards-addenda/addenda-for-standard-34-1997()PDF). ANSI/ASHRAE Standard 34-1997, Designation and Safety Classification of Refrigerants. 1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2000-11-15I.SSN 1041- 2336 (https://www.worldcat.org/issn/1041-2336). Retrieved 2016-10-10. Lay summary (https://ashrae.org/standards-research--technology/standards-addenda/addenda-for-s tandard-34-1997).
4. Wieser, Michael E.; Coplen, Tyler B. (2010-12-12)."Atomic weights of the elements 2009 (IUPAC Technical Report)"(https://www.degruyter.com/view/j/pac.2011.83.issue
-2/pac-rep-10-09-14/pac-rep-10-09-14.xml.)Pure and Applied Chemistry. 83 (2). ISSN 1365-3075 (https://www.worldcat.org/issn/1365-3075). doi:10.1351/PAC-REP- 10-09-14 (https://doi.org/10.1351%2FPAC-REP-10-09-14).
5. Wikipedia:Chemical infobox#References
6. Schoen, J. Andrew, "Listing of Refrigerants"(https://web.archive.org/web/2009031918 4345/http://www.jandrewschoen.com/refchar.tpdf) (PDF), Andy's HVAC/R Web Page, archived from the original (http://www.jandrewschoen.com/refchart.pd)f (PDF) on
2009-03-19, retrieved 2011-12-17
7. "Addenda m, n, o, p, q, r, s, t, u, and v to ANSI/ASHRAE Standard 34-2007"(http://as hrae.org/File%20Library/docLib/Public/20080807_34m_thru_34v_final.pdf()PDF). ANSI/ASHRAE Standard 34-2007, Designation and Safety Classification of Refrigerants. 1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2008-06-26I.SSN 1041- 2336 (https://www.worldcat.org/issn/1041-2336). Retrieved 2011-12-18.
8. "Class I Ozone-depleting Substances"(https://web.archive.org/web/2010121010152 8/http://www.epa.gov/ozone/science/ods/classone.html). Science - Ozone Layer Protection. US EPA. 2007. Archived fromthe original (http://www.epa.gov/Ozone/scie nce/ods/classone.html)on 2010-12-10. Retrieved 2010-12-16.
9. "Global Warming Potentials of ODS Substituet s" (https://web.archive.org/web/201010 16052146/http://www.epa.gov/ozone/geninfo/gwps.html). Science - Ozone Layer Protection. US EPA. 2007. Archived fromthe original (http://www.epa.gov/Ozone/geni nfo/gwps.html) on 2010-10-16. Retrieved 2010-12-16.
10. "Class II Ozone-depleting Substances"(https://web.archive.org/web/2010120903180 3/http://www.epa.gov/ozone/science/ods/classtwo.html). Science - Ozone Layer Protection. US EPA. 2007. Archived fromthe original (http://www.epa.gov/Ozone/scie nce/ods/classtwo.html)on 2010-12-09. Retrieved 2010-12-16.
11. Montreal Protocol on Substances that Deplete the Ozone Layer: 2006 Report of the Rigid and Flexible Foams Technical Options Committee (FTOC); 2006 Assessment(h ttp://ozone.unep.org/Assessment_Panels/TEAP/Reports/FOT C/ftoc_assessment_rep ort06.pdf) (PDF), Nairobi, Kenya: United Nations Environment Programme(UNEP) Ozone Secretariat, March 2007,ISBN 978-92-807-2826-2, retrieved 2010-12-16
12. "Addenda e, f, g, and h to ANSI/ASHRAE Standard 34-2010("http://ashrae.org/File% 20Library/docLib/Public/34_2010_efgh_forPosting.pdf)(PDF). ANSI/ASHRAE Standard 34-2010, Designation and Safety Classification of Refrigeran.ts1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2011-02-03.ISSN 1041-2336 (https://www.worldcat.o rg/issn/1041-2336). Retrieved 2011-12-18.
13. "Addenda d, j, l, m, and t to ANSI/ASHRAE Standard 34-2004("http://ashrae.org/Fil e%20Library/docLib/Public/20070613_342004_d_j_l_m_t_on.lypdf) (PDF). ANSI/ASHRAE Standard 34-2004, Designation and Safety Classification of Refrigerants. 1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2007-03-03I.SSN 1041- 2336 (https://www.worldcat.org/issn/1041-2336). Retrieved 2011-12-18.
14. Ravishankara, A. R.; Daniel, John S.; Portmann, Robert W. (2009-08-27), written at Chemical Sciences Division, Earth System Research Laborator,yNational Oceanic and Atmospheric Administration, 325 Broadwa,yBoulder, CO 80305, USA., "Supporting Online Material for - Nitrous Oxide (N2O): The Dominant Ozone- Depleting Substance Emitted in the 21st Century"(https://www.sciencemag.org/conte nt/suppl/2009/08/27/1176985.DC1/Ravishankara.SOM.pdf()PDF), Science, 1200 New York Avenue NW Washington, DC 20005: American Association for the Advancement of Science, 326 (5949), pp. 123–125,Bibcode:2009Sci...326..123R(htt p://adsabs.harvard.edu/abs/2009Sci...326..123R,)PMID 19713491 (https://www.ncbi. nlm.nih.gov/pubmed/19713491,) doi:10.1126/science.1176985(https://doi.org/10.112 6%2Fscience.1176985), "Methane’s influence on ozone is very dependent on altitude. In the troposphere and lower stratosphere it induces ozone production via "smog chemistry"(10), while at higher altitudes it causes ozone losses [...] The net effect is generally a gain in column ozone due to methane increases, both in the stratospheric column and total column(11). Thus, unlike for N2O, a calculated ODP for methane would likely be negative".
15. Critical point (thermodynamics)#Table of liquid–vapor critical temperature and pressure for selected substances
16. Lide, D. R., ed. (2005).CRC Handbook of Chemistry and Physics(86th ed.). Boca Raton (FL): CRC Press. p. 3.22.ISBN 0-8493-0486-5.
17. Hydrocarbons are assumed to be similar to Methane in terms of Atmospheric Lifetime being 12 ± 3 years
18. Pentane & Isopentane is assumed to have a net GWP over 100-yr to be 4 ± 2
19. Hydrocarbons are assumed to be similar to Methane in terms of Semi-Empirical ODP being < 0(smog chemistry)
20. Composition of Refrigerant Blends(http://www.epa.gov/ozone/snap/refrigerants/refble nd.html), US EPA, retrieved 2011-12-18
21. "Addenda af and ag to ANSI/ASHRAE Standard 34-2007"(http://ashrae.org/File%20L ibrary/docLib/Public/20090824_34_2007_af_ag_final.pdf()PDF). ANSI/ASHRAE Standard 34-2007, Designation and Safety Classification of Refrigeran.ts1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2009-06-25.ISSN 1041-2336 (https://www.worldcat.o rg/issn/1041-2336). Retrieved 2011-12-18.
22. "Addenda i, j, k, l, n, and o to ANSI/ASHRAE Standard 34-2010("http://ashrae.org/Fil e%20Library/docLib/Public/34_ijklno_Final.pdf)(PDF). ANSI/ASHRAE Standard 34- 2010, Designation and Safety Classification of Refrigerant.s1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air- Conditioning Engineers, Inc. 2011-06-30.ISSN 1041-2336 (https://www.worldcat.org/i ssn/1041-2336). Retrieved 2011-12-18.
23. "Addenda a, b, and d to ANSI/ASHRAE Standard 34-2010("http://ashrae.org/File%20 Library/docLib/Public/20100916_34_2010_a_b_d_final.pdf()PDF). ANSI/ASHRAE Standard 34-2010, Designation and Safety Classification of Refrigeran.ts1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2010-07-01.ISSN 1041-2336 (https://www.worldcat.o rg/issn/1041-2336). Retrieved 2011-12-18.
24. "Addenda a, b, c, d, e, f, g, and h to ANSI/ASHRAE Standard 34-2007("http://ashrae. org/File%20Library/docLib/Public/20070706_34_2007a_b_c_d_e_f_g_h.pdf()PDF). ANSI/ASHRAE Standard 34-2007, Designation and Safety Classification of Refrigerants. 1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2007-06-28I.SSN 1041- 2336 (https://www.worldcat.org/issn/1041-2336). Retrieved 2011-12-18.
25. "Addenda x, y, aa, ab, ac, ad, and ae to ANSI/ASHRAE Standard 34-2007"(http://ash rae.org/File%20Library/docLib/Public/20090224_ad342007x_y_aa_ab_ac_ad_ae.pdf) (PDF). ANSI/ASHRAE Standard 34-2007, Designation and Safety Classification of Refrigerants. 1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2009-01-29I.SSN 1041- 2336 (https://www.worldcat.org/issn/1041-2336). Retrieved 2011-12-18.
26. "Addenda i, j, and k to ANSI/ASHRAE Standard 34-2007"(http://ashrae.org/File%20Li brary/docLib/Public/20080618_34_2007_i_j_k_final.pdf()PDF). ANSI/ASHRAE Standard 34-2007, Designation and Safety Classification of Refrigeran.ts1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2008-01-24.ISSN 1041-2336 (https://www.worldcat.o rg/issn/1041-2336). Retrieved 2011-12-18.
27. PRODUCT SAFETY DATA SHEET - PRODUCT: RS-45 (R434A) (http://www.refsols.c om/files/RS-45/RS-45_safety_data.pdf)(PDF), Refrigerant Solutions Limited, 2001- 10-01, retrieved 2011-12-16
28. TRANSITIONING TO LOW-GWP - ALTERNATIVES IN DOMESTIC REFRIGERATION (http://www.epa.gov/ozone/downloads/EPA_HFC_DomRef.pdf) (PDF), US EPA, October 2010, retrieved 2011-12-17
29. "II. ADVANTAGES OF AMMONIA AS A REFRIGERANT"(https://web.archive.org/we b/20110726172415/https://www.iiar.org/aar/aar_green2.cfm). Ammonia: The Natural Refrigerant of Choice (An IIAR Green Paper.)International Institute of Ammonia Refrigeration. Archived fromthe original (https://www.iiar.org/aar/aar_green2.cfm)on 2011-07-26. Retrieved 2010-12-21.
30. "Ammonia - NH3 - Properties" (http://www.engineeringtoolbox.com/ammonia-d_971.h tml). Engineering Toolbox. Retrieved 2011-12-18.
31. "Refrigerants - Environment Properties"(https://www.webcitation.org/6HtctfLPj?url=htt p://www.engineeringtoolbox.com/Refrigerants-Environment-Properties-d_1220.html.) Engineering ToolBox. Archived fromthe original (http://www.engineeringtoolbox.com/ Refrigerants-Environment-Properties-d_1220.html)on 2013-07-05. Retrieved
2011-12-27.
32. Archer, David (2005). "Fate of fossil fuelCO2 in geologic time" (https://geosci.uchicag o.edu/~archer/reprints/arche.r2005.fate_co2.pdf) (PDF). Journal of Geophysical Research. 110 (C9): C09S05.1–6.Bibcode:2005JGRC..11009S05A(http://adsabs.ha rvard.edu/abs/2005JGRC..11009S05A.) doi:10.1029/2004JC002625(https://doi.org/1 0.1029%2F2004JC002625). Retrieved 27 July 2007.
33. "CDC - NIOSH Pocket Guide to Chemical Hazards - Sulfur dioxide("https://www.cdc. gov/niosh/npg/npgd0575.html), National Institute for Occupational Safety and Healt,h Centers for Disease Control and Prevention1600 Clifton Rd. Atlanta, GA 30333, USA, 2010-11-18, retrieved 2012-02-06
34. "Solstice® LBA Technical Brochure ¦ BlowingAgents" (https://www.honeywell-blowing agents.com/?document=solstice-lba-technical-brochure.)www.honeywell- blowingagents.com. Retrieved 2016-07-05.
35. GM First to Market Greenhouse Gas-Friendly Air Conditioning Refrigerant in U.S(.htt ps://media.gm.com/content/media/us/en/news/news_detail.html/content/Pages/news/ us/en/2010/July/0723_refrigerant,) Warren, Michigan, 2010-07-23, retrieved
2010-12-27
36. "ANSI/ASHRAE Addendum c to ANSI/ASHRAE Standard 34-2010("http://ashrae.org/ File%20Library/docLib/Public/34_2010_c_Final.pdf)(PDF). ANSI/ASHRAE Standard 34-2010, Designation and Safety Classification of Refrigerant.sASHRAE. 2011-07- 27. Retrieved 2011-12-18.
37. "Addendum w to ANSI/ASHRAE Standard 34-2007"(http://ashrae.org/File%20Librar y/docLib/Public/20100120_34_2007_w_final.pdf)(PDF). ANSI/ASHRAE Standard 34- 2007, Designation and Safety Classification of Refrigerant.s1791 Tullie Circle NE, Atlanta, GA 30329-2305: American Society of Heating, Refrigerating and Air- Conditioning Engineers, Inc. 2008-11-18.ISSN 1041-2336 (https://www.worldcat.org/i ssn/1041-2336). Retrieved 2011-12-18.
38. "Numbering Scheme for Ozone-Depleting Substances and their Substitutes("http://w ww.epa.gov/ozone/geninfo/numbers.html.) Science - Ozone Layer Protection. US EPA. 2007. Retrieved 2011-12-18.
39. "ANSI/ASHRAE Addendum ak to ANSI/ASHRAE Standard 34-2007("http://ashrae.or g/File%20Library/docLib/Public/20100309_34_2007_ak_final.pdf()PDF). ANSI/ASHRAE Standard 34-2007, Designation and Safety Classification of Refrigerants. 1791 Tullie Circle NE, Atlanta, GA 30329: American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 2010-02-24I.SSN 1041-2336 (http s://www.worldcat.org/issn/1041-2336). Retrieved 2011-12-18.
40. "ANSI/ASHRAE Addenda z, ah, ai, and aj to ANSI/ASHRAE Standard 34-2007("htt p://ashrae.org/File%20Library/docLib/Public/20100309_34_2007_z_ah_ai_aj_final.pd
f) (PDF). ANSI/ASHRAE Standard 34-2007, Designation and Safety Classification of Refrigerants. [American Society of Heating, Refrigerating and Air-Conditioning Engineers, Inc. 1791 Tullie Circle NE, Atlanta, GA 30329. 2010-01-28.ISSN 1041- 2336 (https://www.worldcat.org/issn/1041-2336). Retrieved 2011-12-18.
41. Wade, Leroy G. Jr. (2006). Organic Chemistry(Sixth ed.). Upper Saddle River, New Jersey: Pearson Prentice Hall. p. 279. ISBN 0-13-147871-0. OCLC 56876670 (https:// www.worldcat.org/oclc/56876670).
Notes
1. Blend Calculation using the following information:[1][17][18] 2. Blend Calculation using the following information[2: ][8][10][9][11][19]
External links
Retrieved from "https://en.wikipedia.org/w/index.php?title=List_of_refrigerants&oldid=80088971"1
This page was last edited on 16 September 2017, at 09:25.
Text is available under theCreative Commons Attribution-ShareAlike License; additional terms may apply. By using this site, you agree to theTerms of Use and Privacy Policy. Wikipedia® is a registered trademark of theWikimedia Foundation, Inc,. a non-profit organization.